METEOROLOGICAL

QUALITY ASSURANCE PROJECT PLAN

VOLUME III

B-003-OAQ-AMB-QA-22-Q-R2

PREPARED BY:

Indiana Department of Environmental Management (IDEM)

Office of Air Quality (OAQ)

Air Monitoring Branch (AMB)

Quality Assurance Section (QAS)

Mail Code 61-54-2

100 N. Senate Avenue

Indianapolis, IN 46204-2251

Revision 2

January 1, 2022

QAPP Revision History

| Revision Number | Date | Responsible Party | Description of Change |
| --- | --- | --- | --- |
| 0 | January 1, 2020 | QAS Chief | New QAPP format to replace QA Manual, which served as OAQ AMB QAPP, and was last U.S. EPA approved on March 9, 2018. |
| 1 | January 1, 2021 | QAS Chief | All Sections: Replaced LEADS with DMDS except on SOP titles; Designated all TSOP titles as SOP; Replaced IDEM extranet with IDEM InfoDump; Removed vertical wind direction Section 3: Added annual review for QAPP Section 6.2: Changed AMS Verification/Calibration for WS and WD from annual to 6 months (NCORE/PAMS) and annual, SR and UVR from bi-annual to 6 months (NCORE/PAMS) and annual, and MLH from annual to 6 months; Changed QAS PE Audit for WS and WD from 6 months (NCORE/PAMS) and annual (SLAMS) to annual; Changed QAS PE for P, SR, and UVR from 6 months to annual Section 7.1: Changed verification/calibration assessment method for WS, WD, SR, and UVR from manufacturer check to collocated check. Measured Quality Objectives changed to match QAS Measured Quality Objectives Section 7.2.1: Added qualifications to table 8 Section 20: Siting changed from 3 years to 2 years |
| 2 | January 1, 2022 | QAS Chief | All Sections: Removed MeteoStar; Replaced LEADS with DMDS in SOP titles List of Acronyms – Removed LEADS Section 6.2, Table 4 – Changed MLH verification/calibration from 6 months to annual Section 7.2.1 – Added checking torque Section 11.2 – Added units that are reported for all parameters Section 12.1 – This section was removed since the information is presented elsewhere in the QAPP and it doesn’t apply to chain-of-custody Section 12.2 – This section was removed since the information does not belong here and it is presented in section 19 Section 12.3 – The information presented in this section was moved to section 10 Section 12.4 – The information presented in this section was moved to section 16 Section 14, Table 11 – Reworded AMS Verification/Calibration Frequency. Changed to “If more than 13 months (7 months for PAMS/NCORE) have passed, data collected past the allowed time frame will be assigned a QA qualifier “2”, operational deviation, pending other checks indicating condition of data. Any indication of an issue may invalidate data starting from the last passing check up to the next passing check. The data would then be replaced with the appropriate null data qualifier.” Section 16, Table 13 – Changed SR and UVR calibration frequency from 2 years to 24 months Section 19, step 1 to 16 - some edits to reflect new DMDS Section 20 – removed Annual Data Certification References – Added References at end of QAPP Revisions made do not require new signatures |

List of Acronyms

| Acronym | Meaning |
| --- | --- |
| °C | Degrees Celsius |
| AA | Administrative Assistant |
| AC | Assistant Commissioner |
| AMB | Air Monitoring Branch |
| AMS | Ambient Monitoring Section |
| ANP | Annual Network Plan |
| AQS | Air Quality System |
| ATS | Air Toxics Section |
| BP | Barometric Pressure |
| CAA | Clean Air Act |
| CBSA | Core Based Statistical Area |
| CFR | Code of Federal Regulations |
| CO | Carbon Monoxide |
| CSN | Chemical Speciation Network |
| CTS | Collocated Transfer System |
| DMDS | Data Management and Display System |
| DP | Dew Point |
| GCMS | Gas Chromatography Mass Spectrometry |
| GD | Guidance Documents |
| RH | Relative Humidity |
| IDEM | Indiana Department of Environmental Management |
| INDOT | Indiana Department of Transportation |
| IMPROVE | Interagency Monitoring of Protected Visual Environments Network |
| MLH | Mixing Layer Height |
| MWG | Maximum Wind Gust |
| NAAQS | National Ambient Air Quality Standards |
| NCore | National Core Network |
| NIST | National Institute of Standards and Technology |
| NO | Nitric Oxide |
| NO2 | Nitrogen Dioxide |
| NOy | Reactive Nitrogen Compounds |
| O3 | Ozone |
| OAQ | Office of Air Quality |
| OPS | Office of Program Support |
| OT | Outdoor Temperature |
| P | Precipitation |
| PAMS | Photochemical Assessment Monitoring Station |
| Pb | Lead |

| Acronym | Meaning |
| --- | --- |
| PE | Performance Evaluation |
| PM | Particulate Matter |
| PM1.0 | Particulate matter having an aerodynamic diameter less than or equal to 1.0 µm |
| PM2.5 | Particulate matter having an aerodynamic diameter less than or equal to 2.5 µm |
| PM10 | Particulate matter having an aerodynamic diameter less than or equal to 10 µm |
| PM10c | Particulate matter having an aerodynamic diameter between 2.5 µm and 10 µm |
| PQAO | Primary Quality Assurance Organization |
| PSD | Prevention of Significant Deterioration |
| QA | Quality Assurance |
| QAPP | Quality Assurance Project Plan |
| QAS | Quality Assurance Section |
| QC | Quality Control |
| QMP | Quality Management Plan |
| REQAS | Recycling, Education and Quality Assurance Section |
| RH | Relative Humidity |
| RPM | Revolutions per Minute |
| RTD | Resistance Temperature Detector |
| RWD | Resultant Wind Direction |
| RWS | Resultant Wind Speed |
| SDHD | Standard Deviation of Horizontal Wind Direction |
| SIP | State Implementation Plan |
| SLAMS | State and Local Air Monitoring Stations |
| SO2 | Sulfur Dioxide |
| SPM | Special Purpose Monitoring |
| SOP | Standard Operating Procedure (this includes Technical Standard Operating Procedures) |
| SR | Solar Radiation |
| STN | Special Trends Network |
| T | Temperature |
| TAD | Technical Assistance Documents |
| U.S. EPA | United States Environmental Protection Agency |
| UVC | Ultraviolet Carbon |
| UVR | Ultraviolet Radiation |
| VFC | Virtual File Cabinet |
| VOC | Volatile Organic Compound |
| WS | Wind Speed |

Section 1:	Quality Assurance Project Plan (QAPP) Identification and Approval

Indiana Department of Environmental Management (IDEM) – Office Air Quality (OAQ) – Air Monitoring Branch (AMB) – Quality Assurance Project Plan (QAPP) – Meteorological – Revision 2

This QAPP is designed to provide an overview of the minimum requirements for a quality assurance (QA) and quality control (QC) program for air monitoring networks which conduct meteorological monitoring in the state of Indiana. Requiring monitoring networks to meet these criteria allows the data from all monitoring networks to be consistent, scientifically defensible, and comparable. Meteorological monitoring consists of:

- Barometric pressure (BP)

- Dew point (DP)

- Maximum wind gust (MWG)

- Mixing layer height (MLH)

- Outdoor temperature (OT)

- Precipitation (P)

- Relative humidity (RH)

- Resultant wind direction (RWD)

- Resultant wind speed (RWS)

- Solar radiation (SR)

- Standard deviation of horizontal wind direction (SDHD)

- Ultraviolet radiation (UVR)

- Wind speed (WS)

A QC and QA program encompasses all phases of ambient air sampling and data analysis. These phases include such activities as site selection, monitoring equipment selection, calibration/verification/audit equipment and procedures, sampling procedures, laboratory analysis, data verification/validation, chain of custody, data reporting, precision/accuracy reporting, and meteorological criteria. Prior to the implementation of any ambient monitoring network becoming operational, a working knowledge of this QAPP is necessary by those personnel designated as QC and QA.

Requirements for meteorological monitoring are available in 40 CFR 58 §58.16, U.S.  Environmental Protection Agency (U. S. EPA) Technical Assistance Documents (TAD) and U.S. EPA QA Guidance Documents (GD). Designated QC and QA personnel should maintain a working knowledge of all applicable requirements. All monitoring and QA program requirements must be kept current and accessible.

Document Approval

Meteorological Quality Assurance Project Plan Volume III

Indiana Department of Environmental Management

Office of Air Quality

Air Monitoring Branch

Indianapolis, IN 46219

B-003-OAQ-AMB-QA-22-Q-R2

This QAPP is hereby recommended for approval and commits the Indiana Department of Environmental Management – Air Monitoring Branch to follow the elements described within.

Approval Signatures and Date Signed:

IDEM AMB		Signature: ______________________________ Date: ____________

Chief

IDEM AMB		Signature: ______________________________ Date: ____________

AMS 1 Chief

IDEM AMB		Signature: ______________________________ Date: ____________

AMS 2 Chief

IDEM AMB		Signature: ______________________________ Date: ____________

ATS Chief

IDEM AMB		Signature: ______________________________ Date: ____________

QAS Chief

IDEM OAQ		Signature: ______________________________ Date: ____________

Assistant Commissioner

IDEM OPS		Signature: ______________________________ Date: ____________

QA Staff

U.S. EPA Region 5	Signature: ______________________________ Date: ____________

Section 2:	Table of Contents

Section 1 Quality Assurance Project Plan Identification and Approval7

Section 2 Table of Contents9

Section 3 Distribution/Notification List12

Section 4 Project/Task Organization12

4.1 Staff Roles and Responsibilities12

4.2 AMB Organizational Chart13

4.3 AMB Roles and Responsibilities14

Section 5 Problem Definition/Background16

Section 6 Project/Task Description16

6.1 Overview of Monitored Meteorological Parameters17

6.2 Project Schedule19

6.3 Site Locations22

Section 7 Quality Objectives and Criteria for Measurement Data22

7.1 Measurement Quality Objectives24

7.2 Collocation Transfer System (CTS Method) 24

7.2.1 Wind CTS Verification Audit25

7.2.2 Radiation CTS Audit25

7.3 Siting26

7.3.1 Wind Speed and Direction26

7.3.2 Outdoor Temperature and Relative Humidity26

7.3.3 Barometric Pressure26

7.3.4 Precipitation26

7.3.5 Solar Radiation and Ultraviolet Radiation26

7.3.6 Mixing Layer Height27

Section 8 Training27

Section 9 Documentation and Records27

Section 10 Network Description (or Sampling Process Design) 28

Section 11 Sampling Method Requirements29

11.1 Meteorological Equipment29

11.2 Monitoring Procedures30

11.3 Failed Data Events30

Section 12 Sample Handling and Custody31

Section 13 Analytical Methods31

Section 14 Quality Control Requirements31

Section 15 Instrument/Equipment Testing, Inspection, and Maintenance Requirements31

Section 16 Instrument Calibration and Frequency32

Section 17 Inspection/Acceptance Requirements for Supplies and Consumables33

Section 18 Non-direct Measurements34

Section 19 Data Management34

Section 20 Assessments and Response Actions36

Section 21 Reports to Management38

Section 22 Data Validation and Usability38

Section 23 Validation and Verification Methods39

23.1 DMDS Data Review40

23.2 AMS Verification40

23.3 QAS Validation43

Section 24 Reconciliation with Data Quality Objectives43

Section 25 References44

List of Tables and Figures

Table 1. QAPP Distribution12

Table 2. Duties of Air Monitoring Branch Positions14

Table 3. Meteorological Data Collection Frequency19

Table 4. AMS Verification/Calibration on Meteorological Equipment20

Table 5. QAS PE Audit on Meteorological Equipment20

Table 6. AMS Meteorological Methods Data Assessment Requirements23

Table 7. QAS Meteorological Methods Data Assessment Requirements24

Table 8. Collocated Wind Audits Limits25

Table 9. Documentation and Records28

Table 10. Meteorological Equipment29

Table 11. Measured Quality Objective Checks and Outcomes31

Table 12. Checks and Preventative Maintenance32

Table 13. Instrument Calibration and Frequency33

Table 14. Assessments and Response Actions36

Table 15. Reports to Management38

Table 16. Null and Data Qualifiers40

# Section 3:	Distribution / Notification List

All members of the IDEM OAQ play an important role in the collection, validation, data analysis, assessment, planning, and reporting of air monitoring data. All entities that are part of this primary quality assurance organization (PQAO) are provided electronic copies of this QAPP and must adhere to the elements of the QAPP. Courtesy copies of the QAPP are also provided to those who conduct air monitoring in Indiana under their own PQAO. Table 1 shows how the QAPP is distributed. An official copy of the QAPP is also available on the IDEM air quality web page and the IDEM SharePoint™ QA Library. The QAPP will be reviewed annually, and documentation provided in the QAPP revision history table.

Table 1. QAPP Distribution

| Name | Organization | Phone |
| --- | --- | --- |
| Air Monitoring Branch chief | IDEM/OAQ/AMB | 317-308-3264 |
| Quality Assurance Section chief and staff | IDEM/OAQ/AMB/QAS | 317-308-3257 |
| Ambient Monitoring Section (1 and 2) chief(s) and staff | IDEM/OAQ/AMB/AMS(s) | AMS#1 317-308-3263 AMS#2 317-308-3260 |
| Air Toxics Section chief and staff | IDEM/OAQ/AMB/ATS | 317-308-3248 |
| Office of Program Support Recycling, Education and Quality Assurance Section chief | IDEM/OPS/REQAS | 317-234-6562 |
| Environmental coordinator | Industries conducting air monitoring in Indiana | Contact QAS Chief |
| Environmental coordinator | Consultants conducting air monitoring in Indiana | Contact QAS Chief |
| QA manager | U.S. EPA Region 5 | 312-353-2325 |
| IDEM Quality Management staff | IDEM Office of Program Support | Contact OPS REQAS Chief |

## Section 4:	Project/Task Organization

## 4.1	Staff Roles and Responsibilities

## Key functions and responsibilities in IDEM are:

- OAQ program management: Assistant Commissioner

- AMB management: AMB chief; AMS (1 and 2) chief(s); ATS chief; QAS chief

- Initiate equipment and supplies request: AMS (1 and 2) environmental managers, QAS environmental managers, and ATS environmental chemists with oversight by AMS (1 and 2) chief(s), QAS chief, and ATS chief

- Procurement of AMB equipment and supplies: final approval by AMB chief; tracking by AA

- Air monitoring site selection, maintenance, and operation which includes calibrations, verifications, span, zero, and QC checks: AMS environmental managers and ATS Chemists with oversight by AMS (1 and 2) Chief(s) and ATS chief. As needed assistance for site selection and parameters by OAQ Programs Branch

- Air monitoring data handling, review, verification, and retrieval requests: AMS environmental managers and ATS chemists with oversight by AMS (1 and 2) chief(s) and ATS chief

- Air monitoring network review and project grants: AMB chief; AMS (1 and 2) chief(s)

- QA performance and system audits, site evaluations, and data validation: QAS environmental managers and program coordinator with oversight by QAS Chief

- QA laboratory: Designated QAS environmental manager oversees most of the work performed in the QA laboratory with some assistance from other QAS environmental managers and oversight by QAS chief

- QMP development/updates, QAPP and SOP approval; SOP agency distribution; review, authorization, and management of QA documentation (part 5 of QMP discusses documents and records): OPS

- Programs Branch, Permits Branch, and Compliance and Enforcement Branch: utilize AMB data; see https://www.in.gov/idem/airquality/ for specific duties of these areas

4.2	AMB Organizational Chart

4.3	AMB Roles and Responsibilities

Table 2 lists general duties of the positions within the AMB. The AMS #2 has an Environmental Manager designated as the AQS administrator, whose responsibilities include data submittal into AQS. Also in the AMS #2 is an Environmental Manager designated as the DMDS administrator, whose duties include reviewing and evaluating data outputs as well as setting limits, overseeing programming within DMDS, and coordinating specific work functions of DMDS with the DMDS contractor. The environmental managers/program coordinator listed under the QAS maintain separate equipment from the AMS(s) and the ATS which ensures an independent QA program. However, on occasion the QAS equipment may be used for a QC check but never to calibrate the site instruments. Data is also validated by the QAS once it has been verified by the AMS(s) and the ATS. The QAS maintains the QAPP(s) and has final decision on data validity.

Table 2. Duties of Air Monitoring Branch Positions

| Position | Duties |
| --- | --- |
| Air Monitoring Branch chief | Overall program management; supervises section chiefs and AA; approves the purchase of major equipment; approves all AMB expenditures approves; QAPPs/SOPs; and approves annual certification of data. |
| Ambient Monitoring Section chiefs | Approves and makes sure AMS staff adhere to the QAPPs/SOPs; oversight and direction of all ambient monitoring functions which includes calibrations, verifications, QC checks, data analysis, site location/setup/shutdown, site maintenance, and the development/update of the ANP/5-year network assessment; ensures data meets quality standards; approves annual certification of data; and supervises AMS staff. |
| Air Toxics Section chief | Approves and makes sure ATS staff adhere to the QAPPs/SOPs; oversight and direction of all toxic functions which include laboratory and field GCMS; instrument calibration and sample analysis; ensures data meets QC standards; aids with the update of the ANP/5-year network assessment; approves annual certification of data; and supervises ATS staff. |
| ## Position | ## Duties |
| ## Quality Assurance Section chief | ## Responsible for the creation, maintenance, revisions, and adherence to the QAPPs/SOPs; oversight and direction of all QA functions which include PE/systems audits, meteorological audits, toxic audits, site evaluations, and operation of the QA laboratory; ensures data meets quality standards with authority to make final decision on data validity; approves annual certification of data; and supervises QAS staff. |
| ## Ambient Monitoring Section environmental managers | ## Performs the daily operations that are required for the air monitoring data to be properly collected, analyzed, and verified; performs site/equipment location/setup/maintenance/shutdown and calibrations/verifications/QC checks on air monitoring field equipment; and reviews, writes, and updates SOPs. |
| ## Air Toxics Section environmental chemists | ## Performs the daily operations that are required for the air monitoring data to be properly collected, analyzed, and verified; perform site visits to conduct maintenance on air toxics monitoring equipment; and reviews, writes, and updates SOPs. |
| ## Quality Assurance Section environmental managers | ## Performs PE/systems audits, meteorological audits, toxic audits, and site evaluations; performs maintenance and calibration/certification/verification on equipment; validates data; reviews, writes, and updates QAPPs/SOPs; and will track the completion of corrective actions and determine the success of these actions. |
| ## Quality Assurance Section program coordinator | ## Assists with meteorological audits and site evaluations; distributes, tracks, and validates data; performs audits on the PM clean rooms; reviews, writes, updates; and distributes QAPPs/SOPs; communicates QA work to AMB Chief for bi-weekly report, which includes SOP approval/revision updates; and will track the completion of corrective actions and determine the success of these actions. |
| ## Air Monitoring Branch administrative assistant | ## Produces, enters, and tracks all requisitions for purchases of AMB equipment, supplies, and services in IDEM’s Purchasing System. |

## Section 5:	Problem Definition/Background

In 1970, the Clean Air Act (CAA) was signed into law. This legislation authorized the development of comprehensive federal and state regulations to limit emissions from both stationary (industrial) sources and mobile sources. The NAAQS addressed six pollutants that threatened public health: SO2, NO2, PM, CO, O3, and Pb. The Board of Health was the regulatory agency in Indiana at the time of the CAA being signed into law. Since 1970, changes in types of pollutants monitored and levels have changed. However, the CAA still provides the regulations and framework for the monitoring of criteria pollutants (CO, NO2, O3, SO2, Pb, PM) by state, local, and tribal organizations through the establishment of an Air Quality Monitoring Program.

In 1986, IDEM was created. Its mission is to implement federal and state regulations to protect human health and the environment while allowing the environmentally sound operations of industrial, agricultural, commercial and government activities vital to a prosperous economy. The mission of the OAQ is to assure all Hoosiers ambient air quality meets the NAAQS; provide timely, quality air permits without unnecessary requirements; and to verify compliance with applicable state and federal air pollution laws and regulations. Five branches are part of the OAQ, which includes Programs, Permits, Compliance and Enforcement, Operations, and Air Monitoring. A description and a flowchart of these branches is available in the QMP at  https://ingov.sharepoint.com/sites/IDEMIntranet/SitePages/Standards,-Policies,-and-Mailcodes.aspx

The AMB is divided among four sections (See 4.2. AMB Organizational Chart, above) which includes two site monitoring sections (AMSs), an air toxics laboratory (ATS), and a quality assurance section (QAS). Meteorological data are collected to:

- Support for pollutant data that are monitored to provide evidence NAAQS is being met

- Use in modeling applications implemented by OAQ Programs Branch which help develop, modify, or activate control strategies that prevent or reduce air pollution episodes

- Determine how weather patterns affect pollution trends throughout the state and/or region

This QAPP covers all the meteorological parameters, as stated in Section 1. The QAPP is reviewed annually and updated if needed. Any SOPs associated with this QAPP are updated a minimum of every four years or if the procedures change.

## Section 6:	Project/Task Description

## Meteorological data are collected since it influences the impact of the pollutants. Data collection has been going on for decades. Some of the meteorological data collection is required for specific sites based on network description. The U.S. EPA provides information on which networks are required to collect specific meteorological data. The National Weather Service may provide meteorological data as well; however, there are instances where it benefits to know the atmospheric condition at the location where the pollutant data are collected. Also, meteorological data generated from the AMB may be more suited for the studies since it is of known accuracy, has specific siting criteria, and is consistent from site to site.

# 6.1	Overview of Monitored Meteorological Parameters

Meteorological parameter data is presented on IDEM’s internet website https://www.in.gov/idem/airquality/. Monthly and annual summary reports of parameters collected are available from continuous monitors. The DMDS provides on-line access to Indiana’s continuous air quality monitoring network. It has been available to the public since July of 2007. DMDS offers access to near real-time data from active sites and historic data from recently discontinued continuous air monitoring sites across Indiana. This allows anyone to track meteorological values throughout the day. Below are the different meteorological parameters which are monitored.

## Collected Meteorological Parameters

### Wind Speed (WS)

### Wind speed is measured directly in miles per hour with either an ultrasonic wind unit or a propeller wind unit. Wind speed values are the average of the 5-minute wind speeds for each hour. This average is calculated by adding all the 5-minute wind speed averages measured in one hour and dividing by the number of measurements. Wind speed is collected since it can affect the dispersal and dilution of pollutants.

### Wind Direction (WD)

Wind direction is measured directly with either an ultrasonic wind unit or a propeller wind unit. Although wind direction is measured, it is not reported as is. The values collected are used to calculate resultant wind speed, resultant wind direction, and standard deviation of horizontal wind direction. Wind direction is collected since it can help determine where the pollutant came from and the transport direction.

Maximum Wind Gust (MWG)

Maximum wind gust is measured in miles per hour with either an ultrasonic wind unit or a propeller wind unit. The maximum wind gust is the peak wind speed observed over several seconds during the hour. Knowing the maximum wind gust value can give a good indication of atmospheric stability and the ability of pollutants to be dispersed.

### Outdoor Temperature (OT)

### Outdoor temperature is measured in degrees Fahrenheit using a platinum RTD temperature probe. This is the temperature of the air outside the monitoring site. Outdoor temperature is measured to determine the amount of rise experienced by a buoyant plume. Outdoor temperature also plays a role in chemical processes of air pollutants.

Relative Humidity (RH)

Relative humidity is measured in percent relative humidity using a Vaisala intercap hygrometer probe. Relative humidity is a measure of the moisture present in the air expressed in percent. One-hundred percent relative humidity is totally saturated air. Humidity is the amount of moisture in the air. Using relative humidity along with temperature and pressure can provide a good indication of chemical reactions of air pollutants.

Barometric Pressure (BP)

Barometric pressure is measured in millibars using a pressure sensor. Barometric pressure is the force per unit area (pressure) exerted by the atmosphere as a consequence of gravitational attraction exerted upon the "column" of air lying directly above a specific point. Knowing the barometric pressure can help determine whether pollution levels build up. When the barometric pressure is high, the air is usually still which allows pollution levels to build up. When the barometric pressure is low, the weather is often wet and windy, causing pollution to be dispersed or washed out of the atmosphere by precipitation.

Precipitation (P)

Precipitation is measured in inches (rainfall) using a tipping bucket rain gauge. Precipitation is any form of water (liquid or solid) that falls from a cloud and reaches the ground. Precipitation can reduce pollution levels in the atmosphere, such as O3 and PM. Knowing if a precipitation event occurred can help determine why pollution levels in the atmosphere decreased.

Solar Radiation (SR)

Solar radiation is measured in watts per square meter using a pyranometer. Solar radiation is the total electromagnetic radiation emitted by the sun which is received at the air monitoring site. Knowing the amount of solar radiation can help in determining photochemical processes of air pollutants in the atmosphere as well as turbulence.

Ultraviolet Radiation (UVR)

Ultraviolet radiation is measured in watts per square meter using a radiometer. Knowing the amount of ultraviolet radiation can help in determining photochemical processes of air pollutants in the atmosphere as well as turbulence.

Mixing Layer Height (MLH)

Mixing layer height is measured in meters using a ceilometer. Measuring the mixing layer height helps determine the volume of air into which pollutants and their precursors are emitted, thus impacting near-surface atmospheric pollutant concentrations.

Calculated Meteorological Parameters

Resultant Wind Speed (RWS)

### Resultant wind speed is measured in miles per hour. Resultant wind speed values are obtained by converting the 5-minute wind speeds and directions for the hour into a single hourly vector. Resultant wind speed is the magnitude of this vector.

### Resultant Wind Direction (RWD)

### Resultant wind direction is measured in degrees compass. Resultant wind direction values are obtained by converting the 5-minute wind speeds and directions for the hour into a single hourly vector. Resultant wind direction is the direction of this vector. Wind direction, the direction from which the wind is blowing, is measured to the nearest degree based on a 360 degree compass with 360 degrees being from the North and 180 degrees being from the South. Zero degrees indicates no wind speed was detected for that time frame (calm conditions).

Standard Deviation of Horizontal Wind Direction (SDHD)

Standard deviation of horizontal wind direction is measured in degrees compass. Standard deviation of horizontal wind direction is a measure of the variability of the direction from which the wind is blowing. Zero standard deviation means that there is no variability in the wind direction. A high standard deviation indicates a large variation in the direction from which the wind is blowing.

Dew Point (DP)

Dew point temperature is measured in degrees Fahrenheit. Dew point temperature values represent the temperature at which the ambient air would become saturated, i.e., dew will form. The dew point can be used to predict the formation of dew, frost, fog, etc. Dew point is calculated by using the outdoor temperature and the relative humidity.

6.2	Project Schedule

## The AMB collects and analyzes meteorological data to assist in determining sources of air pollution and predict pollutant episodes. Checks on the equipment are performed. Tables 3, 4, and 5 list information outlining the collection and checks of the various types of meteorological data. Specific information relating to this work are outlined in the AMB SOPs:

- Audit Method for Precipitation Gauge Measurement

- Eppley Precision Spectral Pyranometer Calibration/Maintenance

- Eppley Total Ultraviolet Radiometer Calibration/Maintenance

- Meteorological Barometric Pressure Calibration/Maintenance

- Precision Spectral Pyranometer and Total Ultraviolet Radiometer Collocated Audit

- RM Young 05305-AQ Wind Direction Audit

- RM Young 05305-AQ Wind Speed Calibration/Maintenance

- RM Young 05305-AQ Wind Unit Audit

- RM Young 41372VF Outdoor Temperature Audit

- RM Young 41372VF Relative Humidity Audit

- RM Young 52202E Precipitation Calibration/Maintenance

- RM Young 86004 Ultrasonic Anemometer Audit

- RM Young Outdoor Temperature/Relative Humidity Calibration/Maintenance

- RM Young/Met One Barometric Pressure Sensor Audit Procedures

Table 3. Meteorological Data Collection Frequency

| ## Parameter | ## Sample Frequency |
| --- | --- |
| ## WS | ## Continuous |
| ## RWS | ## Calculated from continuous WS and WD |
| ## RWD | ## Calculated from continuous WS and WD |
| MWG | ## Continuous |

| Parameter | ## Sample Frequency |
| --- | --- |
| SDHD | ## Calculated from continuous WD |
| OT | ## Continuous |
| DP | ## Calculated from continuous OT and RH |
| RH | ## Continuous |
| BP | ## Continuous |
| P | ## Continuous |
| SR | ## Continuous |
| UVR | ## Continuous |
| MLH | ## Continuous |

Table 4. AMS Verification/Calibration on Meteorological Equipment

| Parameter | Frequency |
| --- | --- |
| ## WS | ## Annual (6 months for NCORE/PAMS) |
| ## WD | ## Annual (6 months for NCORE/PAMS) |
| ## OT | ## Annual (6 months for NCORE/PAMS) |
| ## RH | ## Annual (6 months for NCORE/PAMS) |
| ## BP | ## Annual (6 months for NCORE/PAMS) |
| ## P | ## Annual (6 months for NCORE/PAMS) |
| ## SR | ## Annual (6 months for NCORE/PAMS) |
| ## UVR | ## Annual (6 months for NCORE/PAMS) |
| ## MLH | ## Annual |

Table 5. QAS PE Audit on Meteorological Equipment

| ## Parameter | ## Frequency |
| --- | --- |
| ## WS | ## Annual |
| ## WD | ## Annual |
| ## OT | ## Annual |
| ## RH | ## Annual |
| ## BP | ## Annual |
| ## P | ## Annual |
| ## SR | ## Annual |
| ## UVR | ## Annual |
| ## MLH | ## Annual |

## 6.3	Site Locations

## Site locations are available in the IDEM/OAQ/AMB Annual Network Plan and through  https://www.in.gov/idem/airquality/. The locations and meteorological parameters measured will depend on the type of monitoring network. While specific meteorological parameters are required for the NCORE/PAMS and Near-Road Monitoring, additional meteorological parameters are also collected at these monitoring networks as well as other monitoring networks where there is value in having specific meteorological data. Listed below are the different air monitoring networks where meteorological parameters are collected.

### State and Local Air Monitoring Stations (SLAMS)

SLAMS consist of a national network of monitoring sites where the number and distribution of sites is largely determined by the needs of state air pollution authorities.

### Special Purpose Monitoring (SPM)

SPM are designed/intended for use by state and local agencies to collect supportive data for development of SIPs and/or other specific targeted studies such as: point source identification, control strategy effectiveness, etc. If data are used for SIP purposes, SPM sites must meet all federal and state requirements for monitoring methodology and quality assurance.

### National Core Network/Photochemical Assessment Monitoring Station (NCore/PAMS) Monitoring

NCore is a multi-pollutant approach to monitoring. NCore sites are intended to support multiple objectives with a greater emphasis on assessment, research support, and accountability than the traditional SLAMS networks. NCore provides an opportunity to address new directions in monitoring and begin to fill measurement and technological gaps that have accumulated in the networks. Indiana operates one urban NCore site. These sites are required to measure PM2.5, speciated PM2.5, PM10c, O3, SO2, CO, NO, true NO2, NOy, and meteorology. As of June 2019, PAMS is included at NCore sites located in a CBSA with a population of 1,000,000 or more.

### Near-Road Monitoring

On February 9, 2010, the U.S. EPA promulgated monitoring regulations for the NO2 monitoring network. In the new monitoring requirements, state and local air monitoring agencies are required to install near-road NO2 monitoring stations at locations where peak hourly NO2 concentrations are expected to occur within the near-road environment in larger urban areas. Site selection is required to consider traffic volumes, fleet mix, roadway design, traffic congestion patterns, local terrain, and meteorology in determining where a required near-road NO2 monitor should be placed. Indiana operates one near-road monitoring site. IDEM worked with the INDOT to obtain a location for the site. The near-road site is required to measure NO2, CO, and O3. Meteorological parameters are not required by 40 CFR Part 58 for Near-road but is strongly encouraged by U. S. EPA in the near-road TAD. Toxics VOCs and particulates, such as PM2.5, PM1.0, and continuous speciation (black carbon/UVC) are also measured at this site.

# Chemical Speciation Network (CSN)

As part of the PM2.5 NAAQS review completed in 1997, U. S. EPA established a PM2.5 CSN consisting of STN sites and supplemental speciation sites. The CSN is a component of the National PM2.5 Monitoring Network. The goal of the National PM2.5 Monitoring Network is to monitor if the NAAQS are being attained. However, CSN data are not used for attainment or nonattainment decisions but are intended to complement the activities of the larger gravimetric PM2.5 measurement network component. CSN data are used for multiple objectives, including:

- The assessment of data trends;

- The development of effective SIPs and determination of regulatory compliance;

- The development of emission control strategies and tracking progress of control programs;

- Aiding in the interpretation of health studies by linking effects to PM2.5 constituents;

- Characterizing annual and seasonal spatial variation of aerosols; and

- Comparison to chemical speciation data collected from IMPROVE network.

## Section 7:	Quality Objectives and Criteria for Measurement Data

## The primary data quality objective is to ensure that the data collected by the AMB are consistent, of known and adequate quality, supported by adequate calibrations and evaluations, and sufficiently complete to describe the atmospheric state with respect to spatial and temporal distribution. Although no minimum QA requirements are listed in the CFR, the data collected must be shown to be accurate and represent the actual conditions. To determine if the data are accurate, measurement quality objectives based on U.S. EPA document https://www.epa.gov/sites/default/files/2020-10/documents/volume_iv_meteorological_measurements.pdf , the U.S. EPA QAPP “PAMS Required Site Network for Speciated Volatile Organic Compounds, Carbonyls, and Meteorological Parameters Including Mixing Layer Height”, as well as manufacturer recommendations are adopted.

## Collecting quality data begins with properly trained staff and adequate funding to provide the necessary equipment that meets required performance specifications. High quality data also rely on having adequate supplies available, safe monitoring locations that meet U.S. EPA siting requirements, up to date QAPP(s), and SOPs.

## 7.1	Measurement Quality Objectives

## To ensure the quality of the data, verifications and calibrations are performed by the AMS. For some instrumentation this may require sending the unit back to the manufacturer or some other certification facility which can perform the calibration. Verifications are performed by AMS at specific time intervals, when the unit may not be working accurately, or prior to adjusting or swapping out a unit. The verification helps determine the validity of the data by ensuring that the units meet specific limits. Additional PE audits are performed independently by the QAS. The QA results assist with ensuring the unit is working within specifications and can solidify the site operator’s decision on the condition of the data. Tables 6 and 7 list the types of checks as well as limits for each type of meteorological equipment. Calculations are provided in AMB SOPs listed in Section 6.2 of this QAPP.

Table 6. AMS Meteorological Methods Data Assessment Requirements

| Type of Check: Verification/Calibration |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measured Quality Objectives |
| WS | Propeller unit: Check of mph against synchronous motor at minimum of 2 speeds (one within 10 mph and one greater than 20 mph) and a zero check; Ultrasonic unit: collocated check against propeller unit | Propeller unit: <±0.55 mph up to 10.0 mph check, <±5.1% for speeds greater than 10.0 mph, and zero = 0.0; Ultrasonic unit: see table 8 for collocation limits; Ultrasonic sent to manufacturer for wind tunnel check and calibration if unit does not meet collocated check limits |
| WD | Propeller unit: Check against to and from reference point and 360 degrees linearity check; Ultrasonic unit: collocated check against propeller unit | Propeller unit: <±5.1 degrees difference shooting reference point and <10.1 degrees linearity spread; Ultrasonic unit: see table 8 for collocation limits; Ultrasonic sent to manufacturer for wind tunnel check and calibration if unit does not meet collocated check limits |
| OT | Check against temperature transfer standard at approximately 45 °C, 20 °C, 0 °C, and -20 °C | <±1.1 °C (<±0.55 °C for NCORE/PAMS) |
| RH | Check against RH transfer standard at approximately 85%RH, 60%RH, 35%RH, and air check | <±10.1 %RH (<±5.1 %RH for NCORE/PAMS) |
| BP | Check against BP transfer standard 2 times (e.g., arrival at site and prior to departure) | <±3.1 millibars (<±1.1 millibars for NCORE/PAMS) |
| SR | Collocated check; Unit also sent to manufacturer | <±5.1% for daylight values; nighttime values should be 0 (see 7.2.2); Unit sent to manufacturer for check and calibration every 2 years or if fails a collocated check |
| UVR | Collocated check; Unit also sent to manufacturer | <±5.1% for daylight values; nighttime values should be 0 (see 7.2.2); Unit sent to manufacturer for check and calibration every 2 years or if it fails a collocated check |
| P | Check against 3 known amounts of water at approximately 0.295 inches, 0.590 inches, and 0.885 inches | <±10.1% |
| MLH | Check against target placed at a fixed known distance | ≤±5.0 meters or ±1.0%, whichever is greater |

Table 7. QAS Meteorological Methods Data Assessment Requirements

| ## Type of Check: PE Audit |  |  |
| --- | --- | --- |
| ## Parameter Method | ## Assessment Method | ## Measured Quality Objectives |
| WS | Propeller unit: Check of mph against synchronous motor at 4 speeds (approximately 4.58, 11.45, 22.90, and 80.15 mph) and a zero check; Ultrasonic unit: collocated check against propeller unit | Propeller unit: <±0.55 mph up to 10.0 mph check, <±5.1% for speeds greater than 10.0 mph, and zero = 0.0; QAS propeller unit collocated with site ultrasonic (see Table 8) |
| WD | Propeller unit: Check against to and from reference point and 360 degrees linearity check; Ultrasonic unit: collocated check against propeller unit | Propeller unit: <±5.1 degrees difference shooting reference point and <10.1 degrees linearity spread for bench test; QAS propeller unit collocated with site ultrasonic (see Table 8) |
| OT | Check against temperature transfer standard at approximately 45 °C, 20 °C, 0 °C, and -20 °C | <±1.1 °C (<±0.55 °C for NCORE/PAMS) |
| RH | Check against RH transfer standard at approximately 85%RH, 35%RH, and air check | <±10.1 %RH (<±5.1 %RH for NCORE/PAMS) |
| BP | Check against BP transfer standard 2 times (e.g., arrival at site and prior to departure) | <±3.1 millibars (<±1.1 millibars for NCORE/PAMS) |
| SR | Collocated check | <±5.1% for daylight values; nighttime values should be 0 (see 7.2.2) |
| UVR | Collocated check | <±5.1% for daylight values; nighttime values should be 0 (see 7.2.2) |
| P | Check against 2 known amounts of water, one check < 1 inch and check > 1 inch | <±10.1% |
| MLH | Check against target placed at a fixed known distance | ≤±5 meters or ±1%, whichever is greater |

## 7.2	Collocated Transfer System (CTS) Method

## For sites which collect wind data with an ultrasonic wind unit and radiation with a pyranometer and ultraviolet radiometer, verification and audit checks on these systems are performed using a CTS method.

## 7.2.1	Wind CTS Verification and Audit

At sites which have an ultrasonic wind unit, the AMS and QAS will set up a propeller unit for approximately one week to check the wind speed, resultant wind speed, and the resultant wind direction. The propeller unit’s accuracy is confirmed prior to the collocated verification and audit by checking the shaft rotation speed with several known rotation speeds (RPM) and confirming the wind direction passes a bench linearity test. The torque is also tested. The wind speed limits are described in Table 7 of this QAPP, and the bench linearity limit is five degrees. Torque limits are described in the RM Young 05305-AQ manual.

Once installed the CTS wind unit is aligned within three degrees of a known reference point. The degrees of the reference point are determined by using the solar azimuth or a compass (using magnetic declination). The AMB SOP “Determining Degrees to Reference Point using the Solar Azimuth” provides information on determining the degrees. Once the CTS unit is picked up, the data are collected and compared. The comparison is based on taking the mean of all the hourly differences during the test period. The limits are listed in Table 8. Failure to meet the limits will result in screening the data for potential outliers, resolve potential issues with the field or the CTS unit, and overall review of data collected at the site, before, during, and after the check. Any issues with the results may require a QA qualifier or null data qualifier being applied to the data, the field unit needing maintenance/calibration, or additional CTS checks.

Table 8: Collocated Wind Audit Limits

| ## Parameter | ## Average Difference | ## Standard Deviation of difference | ## Qualifications |
| --- | --- | --- | --- |
| ## WS | ## ± 0.55923 mph < 11.185 mph ## ± 5.592 mph > 11.185 mph | ## 0.4474 mph | ## Wind speeds > 2.237 mph |
| ## WD | ## ± 5° | ## 2° | ## Wind speeds > 2.237 mph |

## 7.2.2	Radiation CTS Audit

## At sites which monitor SR and UVR, the AMS and QAS will set up a CTS radiation unit of the same make and model and run it for approximately one week. After the verification or audit is completed, the data are collected and compared. The comparison is based on taking the average difference between values that are above zero. The comparison should be <±5.1% for daylight values. Failure to meet the limits will result in screening the data for potential outliers, resolve potential issues with the field or CTS unit, and overall review of data collected at the site, before, during, and after the check. Any issues with the results may require a QA qualifier or null data qualifier, the field unit needing maintenance/calibration, or additional CTS checks.

## 7.3	Siting

Siting of meteorological equipment for the required measurements is specific to each instrument type. Proper siting is part of the total quality control program. Of course, as in many other monitoring activities, the ideal site may not be attainable. If it is impossible to find sites that meet all siting criteria, compromises will be made, and proper documentation is recorded to reflect this. The important thing to realize is that the data will be compromised, but not necessarily in a random way. The primary objective of meteorological siting is to place the instrument where it can make precise measurements that are representative of the atmosphere in that area. The QAS performs a site evaluation once a site is established, on a 2-year cycle, and if changes are made to the site. Documentation includes sensor heights, obstructions, photographs of the site, and if needed a written description outlining any potential interferences on the sensors. Listed below are specific siting requirements for meteorological equipment.

## 7.3.1	Wind Speed and Direction

The standard height of wind sensors over level, open terrain is 10.0 meters above ground level. If other sensor heights are used, they will be documented accurately in the ANP and AQS. Open terrain is defined as an area where the distance between the sensor and any obstruction is at least 10 times the height of that obstruction. An obstruction may be man-made (e.g. a building) or natural (e.g. a tree). If the sensor is on a tower-mounted boom, the boom must be twice as long as the tower’s diameter and directed into the prevailing wind. If the sensor is mounted up on a tower, the tower must be of open grid design. Sensors mounted on top of a tower will be at least one tower diameter/diagonal above the top of the tower structure. Sensors on a boom will be securely mounted so that they will not twist, rotate, or sway.

## 7.3.2	Outdoor Temperature and Relative Humidity

The standard height is 2.0 meters above the ground. In locations where this is not possible, the sensor will be placed at approximately 10 meters, such as on top of an open grid design tower where the wind sensor is located. For the sensors at 2.0 meters, the ground beneath the sensor must be non-watered short grass or natural earth, level ground, have an area at least 9.0 meters in diameter, and should be a minimum of 30 meters from large, paved areas. For all sensor heights, large industrial heat sources, roof tops, steep slopes, hollows, high vegetation, swamps, snow drifts, standing water, and air exhausts are avoided. If the sensor is mounted on a boom, the boom must be as long as the tower's diameter or diagonal distance. Sensors will have either a downward facing aspirated shield or one that points to true north. The distance between the sensor and any obstruction, man-made (e.g., a building) or natural (e.g., a tree), must be at least four times the height of the obstruction.

## 7.3.3	Barometric Pressure

The sensor will be located in a ventilated shelter about two meters above ground level. Care is taken to not expose the pressure sensor to anything that could interfere with the measurements. The sensor is sensitive to both the atmospheric pressure (weight above the station) and wind pressure.

## 7.3.4	Precipitation

The mouth of the gauge must be mounted horizontally at a minimum of 30 centimeters above ground. It is positioned in a location to avoid water splashing, being covered by snow, or anything overhead that could obstruct precipitation from going into the gauge. In addition, the collector will be heated to properly measure frozen precipitation. If located at ground level, the surface area around the gauge must be natural vegetation or gravel. The gauge must be positioned at least two times the height of any obstruction. If placed in an open area, a wind shield should be used.

## 7.3.5	Solar Radiation and Ultraviolet Radiation

## The sensor will have an unrestricted view of the sky in all directions. Locations where there are obstructions that could cast a shadow or reflect light on the sensor, such as light-colored walls or artificial sources of radiation, are avoided. The elevation to the horizon as viewed from the pyranometer must not exceed five degrees. Sensor height is not critical for pyranometers; consequently, tall platforms or roof tops are typical locations.

## 7.3.6	Mixing Layer Height

## The ceilometer for determining MLH measurements is intended for more macro-scale application than are the surface meteorological measurements. Consequently, the location of the ceilometer need not be associated with any particular PAMS surface site. Factors that will be considered in selecting a site for the MLH monitor include whether the upper-air measurement for the proposed location are likely to provide the necessary data to characterize the meteorological conditions associated with high ozone concentrations, and the extent to which data for the proposed location of the ceilometer may augment an existing upper-air network. The ceilometer must be installed on a level concrete pad or wooden platform suitably located to provide an unobstructed view of the sky. A wide-open location is recommended where there are no tall trees, overhead lines, or antennas nearby. Proximity to powerful radars should also be avoided. Any object in the cone projecting upward created by an angle of 25° from vertical will impede the ability of the ceilometer to properly measure atmospheric backscatter. Common interfering objects would include powerlines, tree branches, tower support guidewires, flagpoles, or similar features which may permanently or transiently present above the ceilometer.

## Section 8:	Training

Formal staff training is scheduled for new employees and periodically update employees’ skills and program operations. Required training consists of QAPP and SOP review, air monitoring equipment manuals and corresponding SOPs, and hands-on training on air monitoring instruments which are related to the required work for specific staff (e.g., collocated wind setup for QA PE audit; OT/RH blower cleaning). Formal staff training is coordinated with the section chiefs, senior level staff, or parameter specialists of the AMS, QAS, and ATS of the AMB on an as needed basis for those persons(s) engaged in the following: operating, calibrating, verifying, validating, and auditing analyzers/samplers; laboratory procedures; field duties; safety; and any other items related to work performed by staff in the AMB. The training for staff is tracked and documented by the individual section chiefs, except for any in-house training pertaining to computer safety, which is documented by the IDEM computer staff but able to be tracked by individual section chiefs. Standard literature references are readily available to all staff members including the Federal Register, manufacturer’s instrument manuals, and quality assurance guideline documents related to the program objectives. Courses and other training are also provided through U.S. EPA and vendors.

## Section 9:	Documentation and Records

## The goal of IDEM is to collect data that are accurate and representative of the actual conditions. For this to occur, documentation and record keeping must be performed at a high level of accuracy and be consistent amongst all participants who are part of the PQAO. Table 9 summarizes what documentation is involved, how it is handled, and how records are maintained.

Table 9. Documentation and Records

| Document | Location | Retention Time | Custodian |
| --- | --- | --- | --- |
| ANP; 5 Year Network Plan; QAPP | IDEM internet and InfoDump; AMB shared drive | Latest on IDEM internet and InfoDump; AMB shared drive maintains previous versions | ANP and 5 Year Network Plan – AMS (1 and 2) Chief(s); QAPP – QAS Chief |
| SOPs | IDEM InfoDump; AMB shared drive | Latest on IDEM InfoDump; AMB shared drive maintains previous SOPs | QAS Program Coordinator and OPS |
| Logs | Log on site computer hard drive; operator log available through DMDS | Indefinitely | AMS Environmental Manager DMDS Administrator |
| AMS Calibration Forms | AMB shared drive | Indefinitely | AMS Parameter Specialist |
| QAS Audit Forms | AMB shared drive | Kept indefinitely | QAS Environmental Manager |
| QAS data memos; data checks; and site evaluations | AMB shared drive; VFC; site evaluation record also on DMDS | Kept indefinitely | QAS Chief and Program Coordinator |
| Calibrations, Certifications, and Verifications performed by the QA laboratory | AMB shared drive | Information kept at least 3 years unless item is still in circulation then information is kept indefinitely | QA Laboratory Manager |
| NIST-traceable Certifications for AMB standard equipment | QA Laboratory Cabinet File | Kept Indefinitely | QA Laboratory Manager |

## Section 10:	Network Description (or Sampling Process Design)

## Meteorological monitoring is primarily conducted based on needs for specific weather patterns that can play a role on pollutant concentrations. The IDEM ANP provides information on sites and can be found at https://www.in.gov/idem/airquality/

## Network design and sampler siting for meteorological measurements is established based “Quality Assurance Handbook for Air Pollution Measurement Systems, Volume IV: Meteorological Measurements", U.S. EPA GD’s, and is mentioned in Sections 6 and 7 of this QAPP.

## When photographs or digital images are taken with either a cell phone or the AMB digital camera for purposes of documenting and to support a field investigation, a record of each exposure or image will be saved as a file on an AMB shared drive on the computer. The following information will be recorded:

## •	The site name and what it shows will be part of the file’s name, which will be stored based on the year it was taken. For example, file name “Gary IITRI SE” stored under the path QA\Site Information\Site Evaluations\Gary IITRI-180890022\2020.

## •	The name of the individual who took the photograph or digital image will correspond to any paperwork, such as a site’s evaluation. If no paperwork is used, a log entry on DMDS is adequate.

## Section 11:	Sampling Method Requirements

## Sampling equipment and procedures for meteorological parameters follow “Quality Assurance Handbook for Air Pollution Measurement Systems, Volume IV: Meteorological Measurements, Version 2.0 (Final)”. Specific instructions on technical aspects of these procedures can be found in the AMB SOPs mentioned in this QAPP and the equipment’s manual. The AMB maintains a complete set of SOPs for all procedures, which is available through the AMB shared computer drive and the IDEM InfoDump, https://ingov.sharepoint.com/sites/IDEMIntranet/SitePages/Standards,-Policies,-and-Mailcodes.aspx

## 11.1	Meteorological Equipment

## IDEM utilizes meteorological equipment that meets accuracy requirements and is able to function for long periods of time with minimal maintenance (see Table 10).

Table 10. Meteorological Equipment

| Parameter | Meteorological Monitoring Equipment |
| --- | --- |
| WS/WD | RM Young model 05305-AQ RM Young model 81000 RM Young model 86004 |
| OT/RH | RM Young model 41372VF |
| BP | Novalynx model 230-601V Novalynx model 230-701 V-5 RM Young model 61202V RM Young model 61302V |
| SR | Eppley model PSP |
| UVR | Eppley model TUVR |
| P | RM Young model 52202-E |
| MLH | Vaisala Model CL51 |

## 11.2	Monitoring Procedures

- WS – A Propeller unit measures the WS based on the revolutions per minute of the propeller. The ultrasonic wind unit operates on the principle that the time required for a sound wave to travel from point A to point B is affected by the speed of the wind in a predictable and repeatable way. Measurements are in miles per hour.

- WD – A Propeller unit measures the WD based on the direction the unit is facing. The wind will move the unit by exerting a force on the tail of the unit. The ultrasonic wind unit operates on the principle that a transducer sends a pulse of ultrasonic sound from the north facing side of the sensor. Measurements are in degrees compass.

- OT – Combination OT/RH probe using platinum RTD located inside a ventilated shield. Measurements are in degrees Fahrenheit.

- RH – Combination OT/RH probe using Vaisala intercap hygrometer located inside a ventilated shield. Measurements are in percent relative humidity.

- BP – A BP sensor measures absolute atmospheric pressure and converts it into a linear, proportional voltage, using digital computer technology. Measurements are in millibars.

- SR – A Precision Spectral Pyranometer is used to measure the solar radiation, which is the total electromagnetic radiation emitted by the sun which is received at our monitoring site. These values are expressed in Watts/m2.

- UVR – A Total Ultraviolet Radiometer is used for the measurement of solar UV radiation. This instrument utilizes a photoelectric cell protected by a quartz window. A specially designed teflon diffuser not only reduces the radiant flux to acceptable levels but also provides close adherence to the Lambert cosine law. An encapsulated narrow bandpass (interference) filter limits the spectral response of the photocell to the wavelength interval 0.295-.0385 µm. Measurements are in Langleys/minute.

- P – A tipping bucket is used for the measurement of precipitation. The tipping bucket consists of a funnel that collects and channels the precipitation into a small seesaw-like container. After a pre-set amount of precipitation falls, the lever tips, dumping the collected water and sending an electrical signal. The tipping bucket is heated so that ice or snow will melt and be reported as rainfall. Measurements are in inches (rainfall)

- MLH – A Ceilometer is designed to measure high-range cirrus cloud heights without surpassing the low and middle layer clouds, or vertical visibility in harsh conditions. It´s backscatter profiling over full range is up to 15 km (49,200 ft.).

## 11.3	Failed Data Events

## In the event of malfunctioning equipment that was not collecting data according to specific requirements, the AMS will document the failure onto the DMDS log, giving detailed information on the condition of the equipment and its effect on the data. If the QAS finds an issue, a detailed log is entered in DMDS, added to the shared drive, and a memo will be sent to the AMS, who will then confirm the results. Refer to the appropriate AMB SOP’s for further information.

## Section 12:	Sample Handling and Custody

# Measurements of all the continuous meteorology are collected in-situ year-round. As no physical samples are collected, chain-of-custody procedures don’t apply.

## Section 13:	Analytical Methods

Analysis at a laboratory is not required because the monitoring method employed analyzes samples in situ.

Section 14:	Quality Control Requirements

Field sampling quality control and acceptance criteria is detailed in Section 7.1 of this QAPP. Table 11 lists the action to take on meteorological data when results do not meet measured quality objectives. Unless proven otherwise, no impact should occur to pollutant data.

Table 11. Measured Quality Objective Checks and Outcomes

| Check | Outcome |
| --- | --- |
| AMS Verification/Calibration Frequency | If more than 13 months (7 months for PAMS/NCORE) have passed, data collected past the allowed time frame will be assigned a QA qualifier “2”, operational deviation, pending other checks indicating condition of data. Any indication of an issue may invalidate data starting from the last passing check up to the next passing check. The data would then be replaced with the appropriate null data qualifier. |
| AMS Verification/Calibration Results | If a verification check fails, further review of the data and the condition of the meteorological equipment may result in either a null data qualifier or a QA Qualifier on the data or void the verification check. A failing verification will require a passing calibration for valid data collection to resume. Also, limits must be met after a calibration/verification for valid data collection to occur. |
| QAS PE Audit Frequency | Any audit that deviates from an assigned check will be made up prior to any AMS verification/calibration. |
| QAS PE Audit Results | If audit results don’t meet specification, data are suspect. The AMS must follow up with a verification, and if needed a calibration. Data may be flagged depending on AMS findings. Discrepancies between AMS and QAS findings may require a repeat of the audit. To come into agreement the QAS may need to observe the AMS verification. |

Section 15:	Instrument and Equipment Testing, Inspection, and Maintenance Requirements

Meteorological equipment must have checks and routine preventive maintenance performed to ensure proper operation. These checks are performed by AMS staff or QAS staff on a routine schedule. Most manufacturers supply a preventive maintenance checklist with the instruction manual. A specific schedule as shown in Table 12 is maintained. All checks and maintenance are documented in DMDS or in an Excel spreadsheet and saved to the AMB shared drive. The parameter specialist is informed of any issues.

# Table 12. Checks and Preventive Maintenance

| Item | Inspection Frequency | Action Required | Action if Item Fails Inspection |
| --- | --- | --- | --- |
| OT/RH Blower | During site check (approximately every 3 weeks). | Determine if blower is working. | Data invalid back to last passing site check unless data review determines otherwise. |
| P | During site check (approximately every 3 weeks). | Clean funnel. | If clogged, data invalid back to last passing check unless data review determines otherwise. |
| SR/UVR | During site check (approximately every 3 weeks). | Ensure sensors are level; check silica gel; ensure no dirt or debris is on top of the units. | Level units; change silica gel; clean top of units; determine if data are accurate or needs a QA Qualifier or a Null Data Qualifier back to last passing check. |
| SR | During site check (approximately every 3 weeks). | Determine if blower is working. | Data suspect back to last passing check unless data review determines otherwise. |
| Wind unit for collocated checks | Prior to doing a collocated verification or audit check. | Determine if unit is working accurately. | Repair or send to manufacturer. |
| Spare meteorological units | When manufacturer calibration date for SR and UV is due. Other met parameters, as needed. | Determine if units are working. Determine if SR and UV meet manufacturer recommended calibration time frame. | Repair or send to manufacturer. |

Section 16:	Instrument Calibration and Frequency

All meteorological equipment in the air monitoring network adheres to the prescribed calibration schedules that are defined in the https://www.epa.gov/sites/default/files/2020-10/documents/volume_iv_meteorological_measurements.pdf  and the AMB SOPs, which are mentioned throughout this QAPP. Transfer standards used for this work also follow strict verification/audit/certification/calibration schedules, as required by the documents mentioned. Almost all transfer standards are verified/certified/calibrated against standards that are maintained at the QA laboratory. If the QA laboratory cannot do the work, then the item is sent to a standards laboratory. On occasion some items are sent to a standards laboratory to cross check the transfer standards used in the QA laboratory. All equipment is verified/certified/calibrated prior to use. Any equipment used in the meteorological program to perform calibrations, verifications, and audits has documentation kept on it. A certification file is kept for each item and stored in the QA laboratory and on the AMB shared drive. Table 13 lists the transfer standards used in the meteorological program.

# Table 13. Instrument Calibration and Frequency

| Type of Device | Frequency | Primary Standard | Limit |
| --- | --- | --- | --- |
| Anemometer Drive or Synchronous Wind Motor | 12 months | QAS NIST-Traceable Photo Tachometer | # ≤±1.0 RPM |
| Barometer | 12 months | QAS Primary Standard Mercury Barometer | # ≤±1.0 millibar |
| Relative Humidity Sensor | 12 months | QAS Manufacturer Certified NIST-Traceable Transfer Standard | # ≤±3.0 %RH |
| Isotech | 12 months | QAS Primary Standard Temperature Device | # ≤+ 0.2 oC |
| Temperature Probe | 12 months | QAS Primary Standard Temperature Device | # ≤+ 0.5 oC; # ≤+ 0.2 oC if used for OT Meteorology Program Calibrations, Verifications, and Audits |
| SR unit used for CTS checks | 24 months | Manufacturer NIST-Traceable Standard | # Limits determined by manufacturer |
| UVR unit used for CTS checks | 24 months | Manufacturer NIST-Traceable Standard | # Limits determined by manufacturer |
| Photo Tachometer | 12 months | Manufacturer NIST-Traceable Standard | # Limits determined by manufacturer |

## Section 17:	Inspection/Acceptance Requirements for Supplies and Consumables

Meteorological replacement parts are obtained from the manufacturer or a distributor when the items are commonly available and needed by the AMB. The parameter specialist for the AMS keeps track of their own supplies and will order items when needed. The QA laboratory manager makes sure the QA laboratory has enough supplies and QA field staff also keep track of when items are needed. A designated QA environmental manager will order supplies for the QAS when needed. Most consumables do not have an expiration date. If there is an expiration date the AMS parameter specialist and the QA laboratory manager will track and dispose of it when it expires.

## Section 18:	Non-direct Measurements

All IDEM meteorological data are produced directly from the IDEM owned NIST-traceable equipment. Some site selection and data analysis can be based upon meteorological data generated by other entities, such as the National Weather Service, as well as traffic counts by INDOT. Census population data to determine the appropriate number of samplers/analyzers for an urban area. Data acquired from non-direct measurements may also include site operator observations. The additional data used for site selection and data analysis are viewed as being accurate since this data fall under specific rules and guidelines.

## Section 19:	Data Management

Minimizing data loss is of paramount importance to each monitoring program in order to meet and exceed the program’s data completeness requirements. Data loss can result from missing or invalid data. Data records are generated using DMDS. It is the goal of IDEM to collect 100% of generated data and maintain a completeness rate of at least 75% valid data. The U.S. EPA requires at least 75% of data to be valid to meet completeness requirements. The processes of determining valid data and the QC/QA of these processes are mentioned in this QAPP and SOPs. The data process includes collecting, storing, transmitting, verifying, validating, and reporting to the U.S. EPA’s AQS database.

## The data process includes collecting, storing, transmitting, verifying, validating, and reporting to the U.S. EPA’s AQS database.

## The data process starting from the field to AQS submittal consists of the following steps below.

- Data collection is performed using meteorological instrumentation of known accuracy and precision. The meteorological instruments are verified/calibrated prior to sample collection. This information is stored on an AMS verification/calibration sheet and logged in DMDS.

- Data are sent from the instrument digitally to the on-site data logger. Although data collection is continuous, data is averaged in five-minute intervals. The data logger can store up to approximately two weeks of data before it starts being overwritten. The data logger requires a password to access it.

- The DMDS communicates with the data loggers and retrieves data every 15 minutes through a cellular modem. The DMDS software collects data from the data logger using the Native Datalogger Computer-To-Computer language (CC-SAIL). Preliminary data checks are performed automatically before data is stored on the DMDS server.

- The data is decoded, checked for errors, processed, and stored in a database. The DMDS serve has a mirror backed for disaster recovery. The raw data can also be reloaded if a needed, such as accidental deletion of processed numbers or erroneous slope, intercept, offset applied to data.

- Each business day, the AMS DMDS Administrator will check the monitoring data database to ensure that all data were polled and transmitted successfully from each monitoring station and stored on DMDS. In case of missing data, a re-poll can be initiated to backfill the data.

- Approximately one to four weeks after the end of a month, AMS staff will initiate the data verification procedure. AMS staff review and flag raw data and document data verification within the DMDS. The DMDS requires a log entry when any changes are made to the data. The logs document who verified the data, changes made, and a date/time when this occurred (see AMB SOP “Meteorological Parameters Data Validation”).

- Once the data have been verified, AMS staff will inform the QAS program coordinator via e-mail that verified data are ready for validation.

- The QAS program coordinator logs the date that the verified data has been sent to the QAS and then informs assigned QAS environmental manager via e-mail that data is ready for validation. The QAS environmental manager has 15 business days to complete the validation process.

- The QAS environmental manager logs into the DMDS server (read access only) and performs an audit on the data (validation process) using the various DMDS reports. The validation process is documented on a Validation Check form, which is stored on the AMB shared drive.

- If data issues arise during the validation process, the issues are sent to the appropriate AMS staff member for correction or additional information. Any data corrections will be noted in the validator logs.

- Once the validation review is complete, the QAS environmental manager will electronically initial and date the Validation Check form. The QAS environmental manager informs the QAS program coordinator via e-mail that the data validation is complete.

- The QAS program coordinator will log the data validation completion date and informs the AMS AQS administrator via email that the data review process is completed by the QAS.

- The AMS AQS administrator downloads the validated data from the DMDS and submits the data to AQS.

- The AMS AQS administrator generates an AQS report on the DMDS server. The AQS report is reviewed for accuracy and submitted to AQS.

## Annually, the AMS will verify each data logger’s channel voltage input against an NIST certified power supply. For sites which use digital, 25% of those sites have the raw data from the data logger compared with the five minute and hourly analog values downloaded from the sensor. Both procedures help ensure an accurate collection of data by testing the acceptability of the hardware and software configurations.

## Quarterly verified and validated data is submitted to the AQS within 90 days after the quarter is complete. Ambient data are made available to the IDEM website as it is being collected; therefore, the data have not initially had QC or QA checks performed. Data are also provided to AIRNOW as it is being collected.

Data generated by the AMS(s) during verifications/calibrations and by QAS during audits is entered into an excel-based form then stored on a shared computer drive.

## Section 20:	Assessments and Response Actions

## Every monitoring project includes periodic assessments to provide continuing verification that the project is being performed according to the procedures developed in the QAPP, and that the data obtained meet the measurement quality and overall monitoring project objectives. This section details the assessments that will be performed, the personnel responsible, the frequency and the reports produced, as well as corrective actions taken based on the results of each assessment. IDEM utilizes several assessment procedures to identify and correct issues. The corrective action process must either include formal communications (e.g., official memoranda) or informal communication (e.g., e-mail messaging) and with responses provided as formal or informal communications. Verbal communication can be used to initiate the corrective action or acknowledge completion of the corrective action but must be followed up with written communication for documentation purposes (see Table 14 for a list of assessments). In any situation, it is the AMB goal to continue to collect valid data; however, there could be some instances where outside factors jeopardize the project, resulting in some data not being collected. The AMB does what it can to maintain continuous data collection as stated in this QAPP and will look for any possible solution to resolve issues.

Table 14. Assessments and Response Actions

| Assessment | Conducted By | Frequency | Goals |
| --- | --- | --- | --- |
| Verification/Calibration | AMS Personnel | Annual or 6 months | QAPP Requirements |
| PE Audit | QAS Personnel | Annual or 6 months | QAPP Requirements |
| Data Verification | AMS | Monthly | Screening of data, chart trace, QC checks |

| Assessment | Conducted By | Frequency | Goals |
| --- | --- | --- | --- |
| Data Validation | QAS | Monthly | Screening of data, chart trace, QC and QA checks |
| QAPP | QAS | Annually | Determine if changes are needed which accurately describes the project |
| ANP | AMB | Annually | Determine if sites cover necessary air monitoring requirements |
| Siting | QAS | 2 years per site | QAPP Requirements |
| 5 Year Network Assessment | AMB | 5 years | Determine future monitoring needs and site locations |
| Technical Systems Audit | U. S. EPA Regional Personnel | Once every 3 years | CFR Requirements |

## In the event an assessment identifies an area of concern, there are specific corrective actions which occur depending on what the finding shows. Below is listed the assessment and corrective action time frame for follow-up.

## Verification/Calibration – The AMS DMDS Administrator will review data generated from the sites daily. The parameter specialist will also review the data several times a week. The AMS addresses issues once they are aware. Verifications/calibrations are performed on a six month or annual basis or when there is evidence of data issues. QAS environmental manager ensures action was taken and issues were resolved.

## PE Audit – The QAS is set up as an independent section within the AMB and as such performs internal PE audits using equipment that are independent of equipment used by the AMS. PE audits are scheduled by a QAS environmental manager on an annual basis. QAS staff members notify the AMS parameter specialist of any issues either from site or once back in the office with written notification (memorandum or e-mail) following soon after. Once the issue is resolved, the AMS parameter specialist must document the issue resolution.

## Data Verification Process – The AMS parameter specialist performs a verification of the ambient data within 1 to 3 weeks after the end of a month. Documentation of the process can be found in the DMDS validator notes.

Data Validation Process – A QAS staff member performs a validation review of the data within 15 working days after the QA Program Coordinator is informed that a data package is available for review. Data validation is documented on the Data Check Sheet, which is stored on the AMB shared drive.

## QAPP – QAS section chief ensures QAPP is being followed and makes necessary changes, with approval by AMB chief, AMS (1 and 2) chiefs, and ATS chief. An annual check is documented or when a change needs to be made.

ANP – The Annual Network Plan is due to the U.S. EPA Regional Administrator by July 1st. The AMB tries to have a complete ANP available for public comment by mid-May to allow for the 30-day public comment period to be completed by mid-June. Corrective actions taken immediately (based on issue could be one day but prior to New Year) by AMB based on U.S. EPA feedback.

## Siting – The QAS performs site evaluations on a 2-year cycle from the previous site evaluation.  The AMS addresses issues based on QAS findings (usually within one workday if data are impacted or some time frames may be extended based on the nature of the issue and if the site is on private property). QAS chief ensures corrective action is taken to resolve the issues.

## Five Year Network Assessment – The 5-year Network Assessment is due to U.S. EPA by July 1st for years ending in zero or five. Corrective actions by AMB based on U.S. EPA feedback within the time frame allotted for the response.

## Technical Systems Audit – Technical System Audits are scheduled by the U.S. EPA Regions on a 3-year frequency. The QAS works with AMS(s) and ATS to address any audit findings. All findings should be resolved within one year of the TSA report.

## Section 21:	Reports to Management

## Reports that are generated and utilized in the meteorological program are listed in Table 15.

Table 15. Reports to Management

| ## Report | ## Frequency | ## Responsible Party |
| --- | --- | --- |
| ## QA Audit Findings | ## As needed | ## QAS |
| ## Invalid Data Memos | ## As needed | ## AMB |
| ## Annual Network Plan | ## Annual | ## AMB |
| ## 5 Year Network Plan | ## 5 Years | ## AMB |

## Section 22:	Data Validation and Usability

Many of the criteria used to review and validate data have been detailed in the previous sections of this QAPP. The AMB utilizes the established QAPP, SOPs, and U.S. EPA document https://www.epa.gov/sites/default/files/2020-10/documents/volume_iv_meteorological_measurements.pdf , to determine data validity and usability. Once data has gone through all the processes, then it is deemed usable for its intended objective.

## Section 23:	Validation and Verification Methods

Data verification is the process of evaluating the completeness, correctness, and conformance of a specific data set against the method, procedural or contractual requirements, as specified in both the SOPs and Volume IV: Meteorological Measurements (EPA-454/B-08-002, March 2008). Data validation is a process that extends the evaluation of data beyond method, procedural or contractual compliance (i.e., data verification) to ensure that reported values meet the quality goals of the environmental data operations and that the data can be used for its intended purpose.

The AMB uses the meteorological validation templates provided in chapter 10 and appendix C of the U.S. EPA QA Handbook for Air Pollution Measurement Systems: Volume IV: Meteorological Measurements (EPA-454/B-08-002, March 2008) as well as manufacturers of the meteorological equipment for the handling of the meteorological data and the weight of evidence approach for validating meteorological data. The grading of the meteorological data is similar to how the pollutant data is handled. The handling of the pollutant data is as follows:

- Critical criteria are issues deemed critical to maintaining the integrity of the hourly ambient concentration measurement or a group of successive hourly ambient concentration measurements. Data reviewers should invalidate observations that do not meet all criterion in the critical criteria table unless there are compelling reasons and justification for not doing so. Basically, the hourly measurement or group of hourly measurements that do not meet one or more of these criteria is invalid unless proven otherwise. In most cases, the CFR dictates the requirement, the implementation frequency of the criteria and the acceptance criteria so these criteria are considered regulatory in nature.

- Operational criteria are situations where violations of a criterion or criteria may be cause for invalidation of the data. Data reviewers should consider other QC information that may or may not indicate the data are acceptable for the parameter they want to control. Therefore, ambient data, which do not meet one or more of these criteria, are suspect unless other QC information demonstrates otherwise, and the reviewers have adequate documentation of that information. Data reviewers should investigate, mitigate, or justify the reason for not meeting the criteria.

- Systematic criteria include those criteria, including the DQOs, which are important for the correct interpretation of the data, but do not usually impact the validity of the ambient data. If the data do not meet the DQOs, this does not invalidate any of the hourly measurements, but it may impact the confidence in the attainment/non-attainment decision.

The AMB brackets all meteorological data using the results of verification checks, calibrations, or PE audits to ensure the meteorological equipment was in proper operating condition between the checks. When equipment fails a check, the operator can flag the data as LIM from the last passing check to time & date when the issue causing failure is remedied. The DMDS program is also able to flag data automatically based on the criteria established in the program, such as extreme outdoor temperature change from one hour to the next. The AMS parameter specialist will review the data, determine the cause for the failure and verify the extent of the data invalidation period. During the validation process, the QAS will review the invalid data period to ensure it is proper, accurate, and documented.

23.1 DMDS Data Review

The DMDS software acquires 5-minutes parameter averages from the instantaneous data generated by the meteorological equipment. As data are being collected in DMDS, there are several automatic checks that the DMDS software performs, such as no changes in data for several hours, extreme values, and stability issues. Dramatic shifts in data can also be recognized by DMDS. The DMDS will flag meteorological data with a LIM flag when the data falls outside the U. S. EPA guidelines for meteorological data collection. The AMS parameter specialist will do a further review of the data and either apply additional flags or adjust any flags applied by the DMDS. DMDS will flag missing data with the DMDS LOST code if it cannot repoll the data logger to download the missing data.

23.2 AMS Verification

The DMDS administrator reviews the DMDS reports every business day to check for anomalies and to repoll data loggers when missing data is notated. In addition, the DMDS administrator will review any LIM flag and notify the AMS parameter specialist of the issue.

After a month of data is collected, a monthly report is generated by DMDS. The AMS verifies all data during the monthly data review per the AMB SOP Meteorological Parameters Data Validation. The monthly report is reviewed by the AMS parameter specialist for data values as well as any flags applied by DMDS. These are reviewed using the DMDS operator to justify the application of the flag and to determine if they are accurately applied. The data trace is also evaluated. Any changes needed are applied by the AMS parameter specialist. Once this process is completed the AMS parameter specialist leaves a “Validator Note” in DMDS, which means the data has been verified. Once the verification process is completed by the AMS, the QAS program coordinator is informed by the AMS parameter specialist that the data is ready for QA.

The null and data qualifiers listed in Table 18 are available to flag data as appropriate. The DMDS will automatically applies a flag text to most. The flag text is then changed to the appropriate qualifier. For data where the DMDS does not automatically apply a flag text, then “NA” is listed under the flag text.

Table 16. Null and Data Qualifiers

| Flag Text | Description | EPA Qualifier |  |
| --- | --- | --- | --- |
|  |  | Null Data | Data Qualifier |
| NA | Deviation from a CFR/Critical Criteria Requirement |  | 1 |
| NA | A 1-Point QC check exceeds acceptance criteria but there is compelling evidence that the analyzer data is valid |  | 1C |
| NA | No 1 Point QC but need to count for completeness |  | 1F |
| NA | Data reviewed and validated |  | 1V |
| NA | Operational Deviation |  | 2 |
| Flag Text | Description | EPA Qualifier |  |
|  |  | Null Data | Data Qualifier |
| NA | Field Issue |  | 3 |
| NA | Lab Issue |  | 4 |
| NA | Outlier |  | 5 |
| NA | QAPP Issue |  | 6 |
| NA | Below Lowest Calibration Level |  | 7 |
| NA | QA/QC Unknown |  | 8 |
| ZERO | Neg Value Detected - Zero Reported |  | 9 |
| AMB-A | High Winds |  | A |
| AMB-B | Stratospheric Ozone Intrusion |  | B |
| AMB-C | Volcanic Eruption |  | C |
| AMB-D | Sandblasting |  | D |
| AMB-E | Forest Fire |  | E |
| AMB-F | Structural Fire |  | F |
| AMB-G | High Pollen Count |  | G |
| AMB-H | Chemical Spill or Industrial Accident |  | H |
| AMB-I | Unusual Traffic Congestion |  | I |
| AMB-J | Construction/Demolition |  | J |
| AMB-K | Agricultural Tilling |  | K |
| AMB-L | Highway Construction |  | L |
| AMB-M | Rerouting of Traffic |  | M |
| AMB-N | Sanding/Salting of Streets |  | N |
| AMB-O | Infrequent Large Gatherings |  | O |
| AMB-P | Roofing Operations |  | P |
| AMB-R | Cleanup After Major Disaster |  | R |
| n | Not used | AA |  |
| TEMP | Shelter Temperature Outside Limits - 9971 - AE | AE |  |
| FEW | Insufficient Data - 9975 - AI | AI |  |
| VOID | Voided by Operator - 9978 - AL | AL |  |
| NEG | Failed NEG Test - 9979 - AM | AM |  |
| MUL | Failed MUL Test - 9979 - AM | AM |  |
| BLAN | Blank Sample | AM |  |
| AUDI | Audit Sample | AM |  |
| Flag Text | Description | EPA Qualifier |  |
|  |  | Null Data | Data Qualifier |
| 5PPB | 5 ppb-V Unblended Standard Check | AM |  |
| EXP | Experimental or Bad Sample | AM |  |
| RT S | Standard Sample | AM |  |
| ARC | GC Acetylene Response Check | AM |  |
| CACS | GC Compress Air Comp Sample | AM |  |
| DCSD | GC Daily Cal Check Stand Dup | AM |  |
| DLA | GC Detection Limit Analysis | AM |  |
| RAS | GC Radian Audit Sample | AM |  |
| UNK | GC Unknown Flag in File | AM |  |
| NOD | Not Detected | AM |  |
| LIM | Failed Limit Check - 9980 - AN | AN |  |
| MAL | Machine Malfunction - 9980 - AN | AN |  |
| ICE | Bad Weather - 9981 - AO | AO |  |
| LOST | Lost - 9983 - AQ | AQ |  |
| POOR | Poor Quality Assurance Results - 9985 - AS | AS |  |
| ADJ | Instrument Adjustment - Cal - Background Zero - 9986- AT | AT |  |
| NOL | Not On Line - 9987 - AU | AU |  |
| POW | Power Failure - 9988 - AV | AV |  |
| SPZ | Span-Zero Check - 9991 - AY | AY |  |
| VER | QC Verification - 9992 - AZ | AZ |  |
| PM | Preventive Maintenance - 9993 - BA | BA |  |
| CAL | Multi-Point Calibration - 9995 - BC | BC |  |
| SPN | Span Check - 9998 - BF | BF |  |
| OPE | Operator Error - 9963 - BJ | BJ |  |
| DAS | Data Logger not communicating with instrument - 9962 - BK | BK |  |
| QAS | QA Audit in Progress - 9961 – BL | BL |  |
| BAL | Negative values under AQS Acceptable Limits - BR | BR |  |
| ALNR | Above Linear Range – EH |  | EH |
| AMB-IF | Fire - Canadian Informational – IF |  | IF |
| AMB-IH | Fireworks Informational – IH |  | IH |
| AMB-IM | Prescribed Burning – IM |  | IM |
| Flag Text | Description | EPA Qualifier |  |
|  |  | Null Data | Data Qualifier |
| AMB-IT | Wildfires – US |  | IT |
| BDL | Below Detection Limit |  | MD |
| AMB-RH | Fireworks RH |  | RH |

Additional flags are available to be applied as well. These can be found at https://aqs.epa.gov/aqsweb/documents/codetables/qualifiers.html.

23.3 QAS Validation

The QAS validates the data per the following AMB SOP “Data Management and Display System (DMDS) Validated Data Review Procedures”. The QAS program coordinator sends an e-mail to the QAS staff member who is responsible for performing the validation on that specific data. A standard form is used by QAS, which includes analysis of data values as well as a review of QC and QA processes and to document the review.

Once data validation is completed by the QAS, the QAS program coordinator informs the AMS AQS administrator that the data validation is complete. The AMS AQS administrator then loads the data into AQS.

The QAS has the authority to have data rechecked and if needed, invalidating additional data.

## Section 24:	Reconciliation with Data Quality Objectives

## The data quality objectives and intended uses for the meteorological data are discussed in Section 7 of this QAPP. The main purpose of this data is to support pollution data which is being collected to show compliance with the U.S. EPA NAAQS. Section 7 of this QAPP also lists the measurement quality objectives, which were established to provide the expected data quality that users need.

## It is the role of the QAPP to establish procedures to control measurement uncertainty to an appropriate level to achieve the objectives for which monitoring data are collected. If guidelines and any SOPs governing the measurement process are followed and all measurement quality objectives listed in this QAPP are met, it will be recognized that the DQOs can be achieved. However, there is always a chance that exceptional field events may negatively affect the performance of the monitoring station. Therefore, it is important to reconcile the monitoring data with the DQOs to evaluate whether the data set is adequate for its intended use. This involves reviewing routine data, such as the monthly verification and validation reviews described in Section 23 of this QAPP, and the results of verification checks. Unacceptable performance for any of the DQO goals does not automatically indicate that the data set cannot be used for its intended purpose, i.e., the support of the decision process for a NAAQS. However, the impact on the confidence with which the data set can be used for its intended purpose in the decision process will have to be reviewed and communicated.

## Section 25:	References

## (1)	EPA QA/R-5 (March 2001), Requirements for Quality Assurance Project Plans, U.S. Environmental Protection Agency, Washington, DC.

## (2)	EPA-454/B-18-006 (August 2018), Guide to Writing Quality Assurance Project Plans for Ambient Air Monitoring Networks, U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Air Quality Assessment Division, Research Triangle Park, NC

## (3)	EPA-454/B-17-001 (January 2017), Quality Assurance Handbook for Air Pollution Measurement Systems, Volume II: Ambient Air Quality Monitoring Program, U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Air Quality Assessment Division, Research Triangle Park, NC

## (4)	EPA-454/B-08-002 (March 2008), Quality Assurance Handbook for Air Pollution Measurement Systems, Volume IV: Meteorological Measurements Version 2.0 (Final), U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Air Quality Analysis Division, Measurement Technology and Ambient Air Monitoring Groups, Research Triangle Park, NC

## (5)	EPA QA/G-8 (November 2002), Guidance on Environmental Data Verification and Data Validation. U.S. Environmental Protection Agency, Washington, DC.

## (6)	EPA-454/B-19-004 (April 2019), Technical Assistance Document for Sampling and Analysis of Ozone, U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Air Quality Assessment Division, Research Triangle Park, NC

## (7)	QAPP: Calibration, Certification, and Verification Methods of Transfer Standards, Volume V (2022)