PARTICULATES

QUALITY ASSURANCE PROJECT PLAN

VOLUME I

B-001-OAQ-AMB-QA-23-Q-R3

PREPARED BY:

Indiana Department of Environmental Management (IDEM)

Office of Air Quality (OAQ)

Air Monitoring Branch (AMB)

Quality Assurance Section (QAS)

Mail Code 61-54-2

100 N. Senate Avenue

Indianapolis, IN 46204-2251

Revision 3

January 1, 2023

QAPP Revision History

| Revision Number | Date | Responsible Party | Description of Change |
| --- | --- | --- | --- |
| 0 | January 1, 2020 | QAS Chief | New QAPP format to replace QA Manual, which served as OAQ AMB QAPP, and was last U.S. EPA approved on March 9, 2018. |
| 1 | January 1, 2021 | QAS Chief | All Sections: Replaced LEADS with DMDS except on SOP titles; Designated all TSOP titles as SOP; Replaced IDEM extranet with IDEM InfoDump; Removed Thermo SHARP and TAPI 602 and any corresponding info. Added Met One 1060 Aethalometer Section 3: Added annual review for QAPP Section 6.2, Table 4 = For TSP for metals intermittent, under sample frequency, added make-up samples may also occur Section 7.1: Removed dust check from QAS audit; Changed PM2.5 Speciation (Aethalometer) passing flow from flow <±15.1% of standard flow to flow <±10.1% of standard flow. Added AT and BP checks to Aethalometer Section 7.3.1: Removed the collocation comparison between the PM2.5 FRM and a continuous PM2.5 FEM sampler performed at a different site each quarter Section 14: Added action to take for late calibration for TSP Section 15: Added information on what action to take when VSCC/SCC and 1st Stage Inlet do not meet time frame for maintenance Section 20: Siting changed from 3 years to 2 years |
| 2 | January 1, 2022 | QAS Chief | All Sections: Removed MeteoStar; Replaced LEADS with DMDS in SOP titles Section 6.2, Table 6 = added minimum of one per month Section 7.1, Tables 12 and 13 = for PM1.0, added if using a T640 for PM1.0, follow the continuous PM2.5 Measured Quality Objectives; for PM10c, removed reference to BAM Section 7.4 = added concentration ranges for Manganese and Arsenic; also added option of adding all three metal solutions to the same strip, with one strip being the low concentration range and one strip being the high concentration range Section 9, Table 14 = added TSA questionnaire, audit finding response form, and final report and closeout letter Section 11.1, Table 15 = added T640 for PM1.0 continuous method; added Met One BC-1060 for PM2.5 speciation continuous; removed BAM for PM10c continuous method Section 12.10 = the information presented in this section was moved to section 10 Section 12.11 = the information presented in this section was moved to section 16 Sections 13.1 and 13.2 = combined into one section; added Microbalance, Anti-static devices, Working standard weights, Primary standard weights Section 14, Table 21 = edited sentence in AMS calibration frequency to include TSP, which is now “data are invalid from 13 months (6 months for TSP) after the calibration until a new calibration is performed” Section 14, Table 22 = added Lab blank ˃ 15 μg, PM2.5 and PM10 field blanks weight gain ˃ 30 μg, Reweigh of filters ˃ 15 μg, and Reporting and collocated samples ˃10.0% Section 15.1 = added solar radiation shield Section 15.4, Table 25 = reworded changing of sample line within 4 years of previous line change unless issues occur Section 19 = some edits to reflect new DMDS Section 21, Table 28 = added Annual Data Certification Package References = added references at end of QAPP Revisions made do not require new signatures |
| 3 | January 1, 2023 | QAS Chief | All Sections = Updated SOP titles; Added and removed some SOPs Section 4.1 = Added AMS program coordinator Section 4.2 = Added AMS program coordinator Section 4.3, Table 2 = Added AMS program coordinator Section 6.2, Table 4 = For PM10 intermittent, added Make-up samples may occur Section 7.1, Table 12 = Removed span dust Section 7.1.1 = Added the following, “If the field blank was properly programmed and ran as a field blank, then it will count towards the 10% collection frequency” Section 7.3.1 = Reworded “In the case of the duplicate sampling method, the collocated sampler must be maintained, operated, calibrated, verified, and audited in the same manner as the reporting sampler” to “In the case of the duplicate sampling method, the collocated sampler must be maintained, operated, calibrated, verified, and audited in the same time frame and in the same manner as the reporting sampler”. Also reworded “CV is calculated only from data pairs (reporting and collocated concentrations) when both values are greater than 3 μg/m3” to “CV is calculated only from data pairs (reporting and collocated concentrations) when both values are equal to or above 3 μg/m3 Section 7.3.2 = Concentration difference percent CV changed from 10 to 20 for when the AMS staff and QAS staff are alerted by the ATS staff Section 11.1, Table 15 = Removed AE21-AR and AE22-ER Section 11.2.2, Table 16 = Removed AE21 and AE22 Section 12.10 = Changed “Staff will limit the number of make-up samples per sampler to 5 per quarter” to “Staff will limit the number of make-up samples per sampler to 5 valid samples per quarter” Section 13.3, Table 17 = Changed checking of working standards against laboratory primary standards from 3 to 6 months to quarterly Section 14, Table 21 = Added “AMS PM2.5, PM10 external leak check, Data suspect back to last passing external leak check. An internal leak check must be performed”. Moved “Filter temperature ˃5 °C above ambient during sampling” from Table 22 to Table 21 Section 14, Table 22 = Added “PM2.5 and PM10 field blanks ran as a sample”. For PM2.5 and PM10 field blanks weight gain ˃ 30 μg, made edit to outcome, added “A QA qualifier “1” is applied to the field blank” Section 15.1, Table 23 = Added PM1.0 - Change the wick every 12 weeks Section 17 = Added TSP to this: All PM2.5, PM10, and TSP filters are acquired through the U.S. EPA national contract Section 19 = Removed (for data reporting) “and some of the continuous speciation, which is at standard conditions of 298 K and 1 atmosphere” Section 19.3 = Added this step “The filters are extracted and analyzed by ATS environmental chemist per AMB SOP “Atomic Absorption Spectrometer Analysis for Lead and other Trace Metals” Section 19.5 = Removed “For some Aethalometer units, the data collected through the processes described in steps 1 to 5 above is overwritten, later, with the data collected on a SD card located inside the Aethalometer. This SD card is collected monthly during AMS staff site visits. Further steps for this data process are available in the AMB SOP “Aethalometer Data Validation” Section 23.4, Table 29 = Deleted Flag Text column; deleted GC flags; and made some edits Section 25 = Added Quality Assurance Guidance Document 2.12 Revisions made do not require new signatures |

List of Acronyms

| Acronym | Meaning |
| --- | --- |
| °C | Degrees Celsius |
| AA | Administrative assistant |
| AC | Assistant commissioner |
| AMB | Air Monitoring Branch |
| AMS | Ambient Monitoring Section |
| ANP | Annual network plan |
| AQI | Air Quality Index |
| AQS | Air Quality System |
| As | Arsenic |
| AT | Ambient temperature |
| ATS | Air Toxics Section |
| BAM | Beta Attenuation Monitor |
| BP | Barometric pressure |
| CAA | Clean Air Act |
| CBSA | Core based statistical area |
| CFR | Code of Federal Regulations |
| CO | Carbon monoxide |
| CV | Coefficient of variation |
| CSN | Chemical Speciation Network |
| DART | Data Analysis Reporting Tool |
| DMDS | Data Management and Display System |
| FEM | Federal Equivalent Method |
| FRM | Federal Reference Method |
| FT | Filter temperature |
| GCMS | Gas chromatography mass spectrometry |
| GD | Guidance documents |
| GLIMS | Gravimetric Laboratory Information Management System |
| IDEM | Indiana Department of Environmental Management |
| INDOT | Indiana Department of Transportation |
| IMPROVE | Interagency Monitoring of Protected Visual Environments Network |
| K | Kelvin; 0 Kelvin is -273.15 °C |
| LPM | Liters per minute |
| Mn | Manganese |
| NAAQS | National Ambient Air Quality Standards |
| NCore | National Core Network |
| NIST | National Institute of Standards and Technology |
| NO | Nitric oxide |
| NO2 | Nitrogen dioxide |

| Acronym | Meaning |
| --- | --- |
| NOy | Reactive nitrogen compounds |
| O3 | Ozone |
| OAQ | Office of Air Quality |
| OPS | Office of Program Support |
| PAMS | Photochemical Assessment Monitoring Station |
| Pb | Lead |
| PE | Performance evaluation |
| PEP | Performance Evaluation Program |
| PM | Particulate matter |
| PM1.0 | Particulate matter having an aerodynamic diameter less than or equal to 1.0 µm |
| PM2.5 | Particulate matter having an aerodynamic diameter less than or equal to 2.5 µm |
| PM10 | Particulate matter having an aerodynamic diameter less than or equal to 10 µm |
| PM10c | Particulate matter having an aerodynamic diameter between 2.5 µm and 10 µm |
| PQAO | Primary quality assurance organization |
| PSD | Prevention of Significant Deterioration |
| QA | Quality assurance |
| QAPP | Quality assurance project plan |
| QAS | Quality Assurance Section |
| QC | Quality control |
| QMP | Quality management plan |
| REQAS | Recycling, Education and Quality Assurance Section |
| RH | Relative humidity |
| SASS | Speciation air sampling system |
| SCC | Sharp cut cyclone |
| SD | Standard deviation |
| SIP | State Implementation Plan |
| SLAMS | State and Local Air Monitoring Stations |
| SO2 | Sulfur dioxide |
| SOP | Standard operating procedure (this includes technical standard operating procedures) |
| SPM | Special purpose monitoring |
| STN | Special Trends Network |
| T | Temperature |
| TAD | Technical assistance documents |
| TAPI | Teledyne Advanced Pollution Instrumentation |
| TEOM | Tapered element oscillating microbalance |
| TSP | Total suspended particulate |
| μg/m3 | Microgram per cubic meter |
| U.S. EPA | United State Environmental Protection Agency |
| URG | University Research Glassware |
| Acronym | Meaning |
| UVC | Ultraviolet carbon |
| VFC | Virtual File Cabinet |
| VOC | Volatile organic compound |
| VSCC | Very sharp cut cyclone |

Section 1:	Quality Assurance Project Plan (QAPP) Identification and Approval

Indiana Department of Environmental Management (IDEM) – Office Air Quality (OAQ) – Air Monitoring Branch (AMB) – Quality Assurance Project Plan (QAPP) – Particulates – Revision 3

This QAPP is designed to provide an overview of the minimum requirements for a quality assurance (QA) and quality control (QC) program for air monitoring networks which conduct particulate sampling in the state of Indiana. Requiring monitoring networks to meet these criteria allows the data from all monitoring networks to be consistent, scientifically defensible, and comparable. Particulate sampling consists of:

- Continuous PM1.0

- Intermittent and continuous PM2.5

- Intermittent and continuous PM10

- Intermittent total suspended particulate (TSP) for metals

- Intermittent and continuous PM2.5 speciation

A QC and QA program encompasses all phases of ambient air sampling and data analysis. The phases include such activities as site selection; monitoring equipment selection; calibration, verification, and audit equipment and procedures; sampling procedures; lab analysis; data verification and validation; chain of custody; data reporting; precision and accuracy reporting; and meteorological criteria. Prior to the implementation of any ambient monitoring network becoming operational, a working knowledge of this QAPP is necessary by those personnel designated as QC and QA.

Three basic sections of the CFR Title 40, Protection of the Environment, which deal with Ambient Air Monitoring are:

- 40 CFR Part 50 which lists the National Primary and Secondary Ambient Air Quality Standards,

- 40 CFR Part 53 which lists alternate equivalent air monitoring methods and procedures for obtaining equivalency, and

- 40 CFR Part 58 which provides detailed descriptions of monitoring methodology, network design and siting, Prevention of Significant Deterioration (PSD) requirements, and QA criteria.

Additional federal requirements are also provided in U.S.  Environmental Protection Agency (U.S. EPA) technical assistance documents (TAD) and U.S. EPA QA guidance documents (GD). Designated QC and QA staff should maintain a working knowledge of all applicable requirements. All monitoring and QA program requirements must be kept current and accessible.

Document Approval

Particulates Quality Assurance Project Plan Volume I

Indiana Department of Environmental Management

Office of Air Quality

Air Monitoring Branch

Indianapolis, IN 46219

B-001-OAQ-AMB-QA-22-Q-R3

This QAPP is hereby recommended for approval and commits the Indiana Department of Environmental Management – Air Monitoring Branch to follow the elements described within.

Approval Signatures and Date Signed:

Section 2:	Table of Contents

Section 1 QA Project Plan (QAPP) Identification and Approval11

Section 2 Table of Contents13

Section 3 Distribution/Notification List17

Section 4 Project/Task Organization17

4.1 Staff Roles and Responsibilities17

4.2 AMB Organizational Chart19

4.3 AMB Roles and Responsibilities19

Section 5 Problem Definition/Background21

Section 6 Project/Task Description22

6.1 Overview of Monitored Particulate Parameters23

6.2 Project Schedule25

6.3 Site Locations28

Section 7 Quality Objectives and Criteria for Measurement Data29

7.1 Measurement Quality Objectives31

7.1.1 Field Blanks34

7.2 PEP Audits35

7.3 Sampler Collocation35

7.3.1 PM2.5/PM1035

7.3.2 Metals36

7.4 Metals Strip Audits37

Section 8 Training38

Section 9 Documentation and Records38

Section 10 Network Description (or Sampling Process Design) 42

Section 11 Sampling Method Requirements43

11.1 Sampling Equipment43

11.2 Sampling Procedures45

11.2.1 Intermittent45

11.2.2 Continuous45

11.3 Failed Sample Events46

Section 12 Sample Handling and Custody47

12.1 Sample Handling47

12.2 Sample Labeling47

12.3 Sample Collection48

12.4 Sample Transportation48

12.5 Sample Custody48

12.6 Continuous Monitoring Data49

12.7 Limited Access49

12.8 Installation and Removal of Data Storage Cards49

12.9 Storage of Raw Data50

12.10 Make-up Samples50

Section 13 Analytical Methods50

13.1 Laboratory Requirements for PM2.5 and PM1050

13.2 Filter Preparation and Analysis for PM2.5 and PM1051

13.3 Sample Weighing for PM2.5 and PM1051

13.4 Environmental Control Requirements for PM2.5 and PM1053

13.5 Permissible Holding Times for PM2.5 and PM1053

13.6 Laboratory Requirements for TSP Filters for Metals54

13.6.1 Laboratory Area54

Section 14 Quality Control Requirements55

Section 15 Instrument/Equipment Testing, Inspection, and Maintenance Requirements60

15.1 PM1.0/PM2.5/PM10 Intermittent and Continuous61

15.2 Metals63

15.3 Intermittent PM2.5 Speciation64

15.3.1 Met One Instruments SASS and Super SASS64

15.3.2 URG 3000N65

15.4 Continuous PM2.5 Speciation65

Section 16 Instrument Calibration and Frequency66

Section 17 Inspection/Acceptance Requirements for Supplies and Consumables67

Section 18 Non-direct Measurements67

Section 19 Data Management68

19.1 Intermittent PM2.5, Intermittent PM10, and Intermittent PMc69

19.2 Intermittent PM2.5 Speciation70

19.3 Intermittent TSP for Metals71

19.4 Continuous PM1.0, Continuous PM2.5, Continuous PM10, and Continuous PMc72

19.5 Continuous PM2.5 Speciation74

Section 20 Assessments and Response Actions75

Section 21 Reports to Management78

Section 22 Data Validation and Usability78

Section 23 Validation and Verification Methods79

23.1 AMS Verification for Intermittent PM2.5, Intermittent PM10, and   Intermittent PMc80

23.2 AMS Verification for Intermittent PM2.5 Speciation80

23.3 AMS Verification for Intermittent TSP for Metals80

23.4 AMS Verification for Continuous PM1, Continuous PM2.5, Continuous PM10, and Continuous PMc80

23.5 AMS Verification for Continuous PM2.5 Speciation83

23.6 QAS Validation83

Section 24 Reconciliation with Data Quality Objectives84

Section 25 References85

List of Tables

Table 1. QAPP Distribution17

Table 2. Duties of Air Monitoring Branch Positions 20

Table 3. NAAQS23

Table 4. Particulate Data Collection Frequency26

Table 5. AMS Calibration Checks on Particulate Samplers26

Table 6. AMS Verification Checks on Particulate Samplers27

Table 7. QAS Audit Checks on Particulate Samplers27

Table 8. Checks on ATS PM Laboratory28

Table 9. Data Quality Indicator/Approach30

Table 10. AMS Intermittent Methods Data Assessment Requirements31

Table 11. QAS Intermittent Methods Data Assessment Requirements32

Table 12. AMS Continuous Methods Data Assessment Requirements33

Table 13. QAS Continuous Methods Data Assessment Requirements34

Table 14. Documentation and Records39

Table 15. PM Sampling Equipment44

Table 16. Continuous Particulate Sampler Indoor Temperature Requirements46

Table 17. Filter Preparation and Analysis Checks51

Table 18. Temperature and Humidity Requirements53

Table 19. Holding Time Restrictions for Filters53

Table 20. Sample Set-up, Run, and Recovery Example Schedule54

Table 21. Measured Quality Objective Checks and Outcomes55

Table 22. Laboratory Measured Quality Objectives and Outcomes58

Table 23. Recommended Maintenance and Frequencies62

Table 24. Inspection of Field Items63

Table 25. Aethalometer Recommended Maintenance66

Table 26. Calibration and Frequency67

Table 27. Assessments and Response Actions75

Table 28. Reports to Management78

Table 29. Null and Data Qualifiers81

# Section 3:	Distribution / Notification List

All IDEM OAQ staff play an important role in the collection, verification, validation, data analysis, assessment, planning, and reporting of air monitoring data. All staff which are part of this primary quality assurance organization (PQAO) are provided with electronic copies of this QAPP and must adhere to the elements of the QAPP. Courtesy copies of the QAPP are also provided to other PQAO conducting air monitoring in Indiana. Table 1 shows to whom the QAPP is distributed. An official copy of the QAPP is also available on the  IDEM air quality web page and the IDEM SharePoint™ QA Library. The QAPP is reviewed annually, and documentation provided in the QAPP revision history table.

Table 1. QAPP Distribution

| Name | Organization | Phone |
| --- | --- | --- |
| Air Monitoring Branch chief | IDEM/OAQ/AMB | 317-308-3264 |
| Quality Assurance Section chief and staff | IDEM/OAQ/AMB/QAS | 317-308-3257 |
| Ambient Monitoring Section (1 and 2) chief(s) and staff | IDEM/OAQ/AMB/AMS(s) | AMS#1 317-308-3263 AMS#2 317-308-3260 |
| Air Toxics Section chief and staff | IDEM/OAQ/AMB/ATS | 317-308-3248 |
| Office of Program Support Recycling, Education and Quality Assurance Section chief | IDEM/OPS/REQAS | 317-234-6562 |
| Environmental coordinator | Industries conducting air monitoring in Indiana | Contact QAS Chief |
| Environmental coordinator | Consultants conducting air monitoring in Indiana | Contact QAS Chief |
| QA manager | U.S. EPA Region 5 | 312-353-2325 |
| IDEM quality management staff | IDEM Office of Program Support | Contact OPS REQAS Chief |

## Section 4:	Project Organization

## 4.1	Staff Roles and Responsibilities

## To comply with the monitoring requirements to determine if areas of Indiana meet the NAAQS and for special air monitoring studies, IDEM is the PQAO for monitoring sites designated as SLAMS, SPM, NCore/PAMS, and Near-Road. The OAQ consists of five branches which includes Air Programs Branch; Permits Branch; Compliance and Enforcement Branch; Billing, Licensing, and Training Branch; and Air Monitoring Branch. A description and a flowchart of the branches is available in the QMP at IDEM InfoDUMP.

## The AMB is divided among four sections (4.2. AMB Organizational Chart) which includes two Ambient Monitoring Sections (AMSs), an Air Toxics Section (ATS) laboratory, and a Quality Assurance Section (QAS). Key functions and responsibilities in IDEM OAQ AMB are:

- OAQ program management: Assistant Commissioner

- AMB management: AMB Chief; AMS (1 and 2) chief(s); ATS chief; QAS chief

- Initiating equipment and supplies requests: AMS (1 and 2) environmental managers, QAS environmental managers, and ATS chemists with oversight by AMS (1 and 2) chief(s), QAS chief, and ATS chief

- Procuring AMB equipment and supplies: Final approval by AMB chief; tracking by AA

- Air monitoring site selection, maintenance, and operation which includes calibrations, verifications, span, zero, and QC checks: AMS environmental managers and ATS chemists with oversight by AMS (1 and 2) chief(s) and ATS chief. As needed, OAQ Programs Branch assists with site selection and parameters

- Air monitoring data handling, review, verification, and retrieval requests: AMS environmental managers and program coordinator and ATS environmental chemists with oversight by AMS (1 and 2) chief(s) and ATS chief

- Air monitoring network review and project grants: AMB chief; AMS (1 and 2) chief(s)

- QA performance and system audits, site evaluations, data validation, and audits of data quality: QAS environmental managers and program coordinator with oversight by QAS chief

- QA laboratory: Designated QAS environmental manager oversees most of the work performed in the QA laboratory with some assistance from other QAS environmental managers and oversight by QAS chief

- Quality management plan development and updates, QAPP and standard operating procedure (SOP) approval; SOP agency distribution; review, authorization, and management of QA documentation (part 5 of QMP discusses documents and records): OPS QA staff

- Programs Branch, Permits Branch, and Compliance and Enforcement Branch: utilize AMB data; see IDEM Air Quality in Indiana for specific duties of these areas

4.2	AMB Organizational Chart

4.3	AMB Roles and Responsibilities

Table 2 lists general duties of the positions within the AMB. An Air Monitoring Section 1 Program Coordinator is designated as the PM filter administrator, whose responsibilities include receiving PM filters from the field staff, logging the filters to the DMDS, and shipping filters to regional staff.  An Air Monitoring Section 2 environmental manager is designated as the air quality system (AQS) administrator, whose responsibilities include data submittal and AMS flow rate verification data being uploaded into AQS. Also, in the AMS #2, an environmental manager is designated as the data management and display system (DMDS) administrator, whose duties include reviewing and evaluating data outputs, setting limits, overseeing programming within DMDS, and coordinating specific work functions of DMDS. A QAS environmental manager is designated to upload the QA flow rate audits into AQS. The environmental managers and program coordinator listed under the QAS maintain separate equipment from the AMS(s) and the ATS ensuring an independent QA program. However, on occasion the same equipment is used for a flow rate verification and a flow rate audit, but the equipment used to calibrate is always independent from the audit. Data are validated by the QAS staff once verified by the AMS(s) and the ATS. The QAS maintains the QAPP(s) and have final decision on data validity.

Table 2. Duties of Air Monitoring Branch Positions

| Position | Duties |
| --- | --- |
| Air Monitoring Branch chief | Overall program management. Approves all branch expenditures. Supervises section chiefs and AA. Approves the purchase of major equipment. Approves QAPPs, SOPs, and annual certification of data. |
| Ambient Monitoring Section chiefs | Approves and ensures AMS staff adhere to the QAPPs and SOPs. Oversee and direct all ambient monitoring functions which includes calibrations; verifications; QC checks; data analysis; site location, setup, and shutdown; site maintenance; and the development or update of the ANP and 5-year network assessment. Ensures data meets quality standards. Approves annual certification of data and supervises AMS staff. |
| Air Toxics Section chief | Approves and ensures ATS staff adhere to the QAPPs and SOPs. Oversees and directs all toxic functions which include laboratory and field GCMS, instrument calibration and sample analysis. Ensures data meets QC standards. Assists with the update of the ANP and 5-year network assessment. Approves annual certification of data and supervises ATS staff. |
| Quality Assurance Section chief | Responsible for the creation, maintenance, revisions, and adherence to the QAPPs and SOPs. Oversees and directs all QA functions which include PE and systems audits, meteorological audits, toxics audits, site evaluations, and operation of the QA laboratory. Ensures data meets quality standards with authority to make final decision on data validity. Tracks the completion of corrective actions and determines the success of these actions. Approves annual certification of data and supervises QAS staff. |
| Air Toxics Section chemists | Performs the daily operations required for the air monitoring data to be properly collected, analyzed, and verified. Performs site visits to conduct maintenance on air toxics monitoring equipment; and reviews, writes, and updates SOPs. |
| Ambient Monitoring Section environmental managers | Performs the daily operations required for the air monitoring data to be properly collected, analyzed, and verified. Performs site and equipment location setup, maintenance, shutdown. Performs calibrations, verifications, and QC checks on air monitoring field equipment; and reviews, writes, and updates SOPs. |

| Position | Duties |
| --- | --- |
| Ambient Monitoring Section program coordinator | Assists with particulate monitoring chain-of-custody through logging of field-collected filters and delivery of filters to the ATS. Ships and receives filters from regional staff. Reviews, updates, and distributes QAPPs and SOPs. |
| Quality Assurance Section environmental managers | Performs PE and systems audits, meteorological audits, toxics audits, and site evaluations. Performs maintenance, calibrations, certifications, and verifications on equipment. Validates data. Reviews, writes, and updates QAPPs and SOPs; and tracks the completion of corrective actions and determines the success of these actions. |
| Quality Assurance Section program coordinator | Assists with meteorological audits and site evaluations. Distributes, tracks, and validates data. Performs audits on the PM clean rooms. Reviews, writes, updates, and distributes QAPPs and SOPs. Communicates QA work to AMB chief for bi-weekly report, which includes SOP approval and revision updates; and tracks the completion of corrective actions and determine the success of these actions. |
| Air Monitoring Branch administrative assistant | Produces, enters, and tracks all requisitions for purchases of AMB equipment, supplies, and services in IDEM’s purchasing system |

## Section 5:	Problem Definition and Background

In 1970, the Clean Air Act (CAA) was signed into law. This legislation authorized the development of comprehensive federal and state regulations to limit emissions from both stationary (industrial) sources and mobile sources. The NAAQS addressed six pollutants that threatened public health: SO2, NO2, PM, CO, O3, and Pb. The Board of Health was the regulatory agency in Indiana at the time of the CAA being signed into law. Since 1970, changes in types of pollutants monitored and levels have changed. However, the CAA still provides the regulations and framework for the monitoring of criteria pollutants (CO, NO2, O3, SO2, Pb, PM) by state, local, and tribal organizations through the establishment of an Air Quality Monitoring Program.

In 1986, IDEM was created. Its mission is to implement federal and state regulations to protect human health and the environment while allowing the environmentally sound operations of industrial, agricultural, commercial and government activities vital to a prosperous economy. The mission of the OAQ is to assure all Hoosiers’ ambient air quality meets the National Ambient Air Quality Standards (NAAQS); provide timely, quality air permits without unnecessary requirements; and to verify compliance with applicable state and federal air pollution laws and regulations.

Air monitoring data are collected to:

- Demonstrate that the NAAQS are being met

- Develop, modify, or activate control strategies which prevent or reduce air pollution episodes

- Detect and analyze pollution trends throughout the state or region

- Provide a database for research and evaluation of effects

This QAPP covers all particulate parameters, as stated in Section 1. The QAPP is reviewed annually and updated when needed. The QAPP is sent to EPA Region V for approval at least every five years. Any SOPs associated with this QAPP are updated a minimum of 4 years or when a procedure changes.

## Section 6:	Project/Task Description

Air quality is regulated to protect public health and the environment in the state of Indiana for decades. The on-going requirement to collect air monitoring data are required by regulation. Data are used to determine compliance with the U.S. EPA’s NAAQS. NAAQS are identified for the criteria pollutants; CO, NO2, O3, SO2, PM2.5, PM10, and Pb. Indiana monitors PM2.5, PM10, and Pb each of which have NAAQS identified (Table 3). Other particulates in this QAPP which do not have established ambient standards are also monitored, including non-criteria metals, PM1.0, and PM2.5 speciated compounds.

Measuring pollutant concentrations in outdoor air and comparing the measured concentrations (e.g., design values) to the corresponding standards determines whether the area’s ambient air quality status is attainment or nonattainment. The NAAQS are separated into primary and secondary standards. Primary standards are established to protect public health. Secondary standards are established to protect the public welfare from adverse pollution effects on soils, water, vegetation, manmade materials, animals, weather, visibility, property, and economy.

The scientific criteria upon which the standards are based are reviewed periodically by the U.S. EPA, which may retain or change the standards according to the findings. [Note hundreds of compounds are considered pollutants when found in ambient air but whose health and welfare effects are not well enough understood for ambient standards to be defined.]

A pollutant measurement which is greater than the ambient air quality standard for its specific averaging time and level is called an exceedance. An exceedance is not necessarily a synonym for a violation. For each pollutant, specific rules regarding the number of allowable exceedances (or concentration above the level) within a given period. In the event the exceedances qualify as a violation of the NAAQS criteria, regulatory actions may result to further clean up the area’s air. The distinction between one exceedance and exceedances which result in a violation allows leeway in the NAAQS for issues caused by unusual weather patterns or unforeseen circumstances.

The design value for a site is the pollutant concentration calculated when the NAAQS rules from 40 CFR Part 50 are applied to the specific pollutant. For example, the PM2.5 design value is calculated by taking the 3-year average of the annual mean. If the number is above the NAAQS for PM2.5, a violation of the standard resulting in the area’s status becoming nonattainment. If the design value is below the NAAQS, then the area’s status is attainment per the standard. Generally, nonattainment is based on the highest design value reported for a specific geographic area, usually a CBSA. If a specific monitor records a violation for the CBSA, the entire area is classified accordingly.

Other important uses of the air monitoring data include the production of a daily Air Quality Index (AQI) report, daily air quality forecast report, support of short and long-term health risk assessments, identification of a localized health concern, and tracking long-term trends in air quality.

## Table 3. NAAQS

| ## Pollutant |  | ## Primary or Secondary | ## Averaging Time | ## Level | ## Form |
| --- | --- | --- | --- | --- | --- |
| ## Particle Pollution | ## PM2.5 | ## Primary | ## 1 Year | ## 12.0 μg/m3 | ## Annual mean, averaged over 3 years |
|  |  | ## Secondary | ## 1 Year | ## 15.0 μg/m3 | ## Annual mean, averaged over 3 years |
|  |  | ## Primary and Secondary | ## 24 Hours | ## 35 μg/m3 | ## 98th percentile, averaged over 3 years |
|  | ## PM10 | ## Primary and Secondary | ## 24 Hours | ## 150 μg/m3 | ## Not to be exceeded more than once per year on average over 3 years |
|  | ## Pb | ## Primary and Secondary | ## Rolling 3-month average | ## 0.15 μg/m3 | ## Not to be exceeded |

# 6.1	Overview of Monitored Particulate Parameters

## Particulate parameter data are presented on IDEM’s internet website IDEM Air Quality in Indiana. Monthly and annual summary reports of pollutants collected by manual methods are available as well as hourly values from continuous monitors. The DMDS provides on-line access to Indiana’s continuous air quality monitoring network. It has been available to the public since July of 2007. DMDS offers access to near real-time data from active sites and historic data from recently discontinued continuous air monitoring sites across Indiana. This allows anyone to track pollutant and meteorological values throughout the day. The following are different particulate parameters which are monitored.

## Criteria Pollutants

### Lead (Pb)

Lead (Pb) is a metal which is highly toxic when ingested or inhaled. It is a suspected carcinogen of the lungs and kidneys and has adverse effects on cardiovascular, nervous, and renal systems.

### Particulate Matter (PM10)

Particulate matter with a mean aerodynamic diameter of 10 microns or less (PM10) is emitted from transportation and industrial sources. Exposure to particle pollution is linked to a variety of significant health problems ranging from aggravated asthma to premature death in people with heart and lung disease. PM10 and PM2.5 data can be used to calculate PM10c.

### Fine Particulate Matter (PM2.5)

Particulate matter with an aerodynamic diameter of 2.5 microns or less (PM2.5) is created primarily from industrial processes and fuel combustion. During breathing these particles can penetrate deeply into the lungs. Exposure to particle pollution is linked to a variety of significant health problems ranging from aggravated asthma to premature death in people with heart and lung disease.

## Non-Criteria Parameters

### PM2.5 Speciation

The U.S. EPA implemented the PM2.5 chemical speciation monitoring program to understand the chemical composition of PM2.5. This information is important for determining sources of pollution and links between observed health effects. The basic objective of speciation analysis is to develop seasonal and annual chemical characterizations of ambient particulates across the nation. This speciation data will be used to perform source attribution analyses, evaluate emission inventories, and air quality models, support health related research studies, and regional haze assessments.

Speciation samplers use different inlets and filters to collect the components of the PM2.5 mixture. One intermittent process consists of using three different types of filters to separate  such specific compounds as: sulfate, nitrate, organic and elemental carbon, ammonium, metals, and certain ions. A continuous process utilizes an instrument which analyzes for black carbon and UVC.

Metals

Metals are categorized as toxic air pollutants, also known as hazardous air pollutants, which are pollutants known or suspected to cause cancer, other serious health effects, or adverse environmental conditions. Examples of non-criteria metals include Manganese (Mn) and Arsenic (As). Chromium is also sampled but due to the very sporadic sampling schedule, this QAPP does not address it in detail. More detailed information on chromium can be found in U.S. EPA SOP(s). Other metals are also analyzed by a contract laboratory when using a Thermo Scientific 2025 for special purpose metals monitoring.

### Ultrafine Particulate Matter (PM1.0)

Particulate matter with an aerodynamic diameter of 1.0 microns or less (PM1.0) is created primarily from industrial processes and fuel combustion. During breathing the particles can deeply penetrate the lungs. Exposure to particle pollution is linked to a variety of significant health problems ranging from aggravated asthma to premature death in people with heart and lung disease.

6.2	Project Schedule

The AMB collects and analyzes particulate samples to determine their concentration except for the intermittent PM2.5 speciation filters. Speciation filters are analyzed by a U.S. EPA contract laboratory. Table 4 lists sampling frequency of various types of particulate data collection. Tables 5 through 8 list specific checks performed on the samplers and the PM laboratory equipment. In-depth details of the laboratory activities are also provided in section 13. Section 20 provides greater detail on non-audit assessments. Specific duties, relating to the work, are outlined in the following AMB SOPs:

- 2025 A/B Sequential Air Sampler Verification and Maintenance

- 2025 A/B Sequential Air Sampler Calibration and Maintenance

- Aethalometer Quarterly Field Audit

- Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software

- Atomic Absorption Spectrometer Analysis for Lead and Other Trace Metals

- Met One Instruments Beta Attenuation Monitor (BAM) 1020 Audit

- Met One Instruments Beta Attenuation Monitor (BAM) 1020 Calibration and Maintenance

- Met One Instruments Beta Attenuation Monitor (BAM) 1020 Verification and Maintenance

- Met One Instruments Speciation Air Sampling (SASS) Calibration and Maintenance

- Met One Instruments Speciation Air Sampling (SASS) Verification and Maintenance

- Met One SASS or SuperSASS PM2.5 Chemical Speciation Sampler – Quarterly Field Audit

- Teledyne API 633 Aethalometer Calibration and Maintenance

- TAPI T640 Calibration and Maintenance

- Thermo Fisher Scientific 2025i Calibration and Maintenance

- Thermo Fisher Tapered Element Oscillating Microbalance (TEOM) 1405 Calibration and Maintenance

- Thermo Scientific Partisol-Plus Model 2025 Audit Procedures

- Thermo Scientific Tapered Element Oscillating Microbalance (TEOM) Quarterly Field Audit

- Tisch Environmental Hi-Vol+ Air Sampler Calibration and Maintenance

- Tisch Environmental Hi-Vol+ Lead Sampler Verification and Maintenance

- Tisch Total Suspended Particulate (TSP) Sampler Audit

- URG 3000N Calibration and Maintenance

- URG 3000N Sequential Particulate Speciation System Quarterly Field Audit

- URG 3000N Verification and Maintenance

Table 4. Particulate Data Collection Frequency

| Sample | Sample Frequency |
| --- | --- |
| PM1.0 continuous | Continuous |
| PM2.5 intermittent | Once every third day except for collocated samplers, which run every sixth day, based on the U.S. EPA sampling schedule. Field blanks are run at least 10% of the time. Make-up samples may occur. |
| PM2.5 continuous | Continuous |
| PM10 intermittent | Once every sixth day. Some 1 in 3 sampling used with PM2.5 to calculate PM10c. Field blanks are run at least 10% of the time. Make-up samples may occur. |
| PM10 continuous | Continuous |
| TSP for metals intermittent | Once every sixth day. Some special study sites will sample extra days, e.g., every third day. Make-up samples may occur. |
| PM2.5 speciation continuous – black carbon/UVC | Continuous |
| PM2.5 speciation intermittent | Once every sixth day. One Super SASS and one URG sampler at the NCORE/PAMS site samples every third day. Field blanks are monthly and does vary from every 4-6 sample runs on a 1 in 6 frequency and every 7-14 sample runs on a 1 in 3 frequency, based on the U.S. EPA speciation schedule. |

Table 5. AMS Calibration Checks on Particulate Samplers

| Parameter | Frequency |
| --- | --- |
| PM1.0 continuous | 365 days – 1 per year with no more than 13 months between two calibrations |
| PM2.5 intermittent | 365 days – 1 per year with no more than 13 months between two calibrations |
| PM2.5 continuous | 365 days – 1 per year with no more than 13 months between two calibrations |
| PM10 intermittent | 365 days – 1 per year with no more than 13 months between two calibrations |
| ## PM10 continuous | ## 365 days – 1 per year with no more than 13 months between two calibrations |
| ## TSP for metals intermittent | ## 90 days |
| ## PM2.5 speciation intermittent | ## 365 days – 1 per year with no more than 13 months between two calibrations |
| ## PM2.5 speciation continuous | ## 365 days – 1 per year with no more than 13 months between two calibrations |

Table 6. AMS Verification Checks on Particulate Samplers

| Parameter | Frequency |
| --- | --- |
| PM1.0 continuous | 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| PM2.5 intermittent | 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| ## PM2.5 continuous | ## 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| ## PM10 intermittent | ## 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| ## PM10 continuous | ## 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| ## TSP for metals intermittent | ## 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| ## PM2.5 speciation intermittent | ## 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |
| ## PM2.5 speciation continuous | ## 30 days – with no less than 21 days and no more than 36 days between verifications and a minimum of one per month |

Table 7. QAS Audit Checks on Particulate Samplers

| Parameter | Frequency |
| --- | --- |
| PM1.0 continuous | 90 days – 5 to 7 months between every other audit |
| PM2.5 intermittent | 90 days – 5 to 7 months between every other audit |
| PM2.5 continuous | 90 days – 5 to 7 months between every other audit |
| ## PM10 intermittent | ## 90 days – 5 to 7 months between every other audit |
| ## PM10 continuous | ## 90 days – 5 to 7 months between every other audit |
| ## TSP for metals intermittent | ## 90 days – 5 to 7 months between every other audit |
| ## PM2.5 speciation intermittent | ## 90 days – 5 to 7 months between every other audit |
| ## PM2.5 speciation continuous | ## 90 days – 5 to 7 months between every other audit |

Table 8. Checks on ATS PM Laboratory

| ## Item | ## Check | ## Frequency |
| --- | --- | --- |
| ## ATS chemists | ## Weighing comparison | ## Annual |
| ## Clean room | ## Internal audit by QAS staff | ## Annual |
| ## Clean room | ## T and RH sensors by QAS staff | ## Quarterly |
| ## PM filter cassettes | ## Leak Check by QAS staff | ## Monthly |
| ## PM filters | ## Re-weighs by QAS staff | ## When filters (clean/exposed) are weighed |
| ## Mass reference standards weights | ## Working vs. primary by ATS staff and QAS staff | ## Quarterly |
| ## Transport and storage cooler temperature | ## Min/Max thermometers by QAS staff | ## Annual |

## 6.3	Site Locations

## Monitoring site locations are available in the IDEM OAQ AMB annual network plan (ANP) and through  IDEM: Air Monitoring: Indiana's Ambient Air Monitoring Network. The locations and particulate parameters measured depends upon the type of monitoring network. Listed below are the different air monitoring networks where particulate parameters are collected.

### State and Local Air Monitoring Stations (SLAMS)

SLAMS consist of a national network of monitoring sites where the number and distribution of sites is largely determined by the needs of state or local air pollution authorities.

### Special Purpose Monitoring (SPM)

SPM is designed and intended for use by state and local agencies to collect supportive data for development of State Implementation Plans (SIPs) or other specific targeted studies such as: point source identification, control strategy effectiveness, etc. If data are used for SIP purposes, SPM sites must meet all federal and state requirements for monitoring methodology and quality assurance.

### National Core Network and Photochemical Assessment Monitoring Station (NCore/PAMS) Monitoring

NCore is a multi-pollutant approach to monitoring. NCore sites are intended to support multiple objectives with a greater emphasis on assessment, research support, and accountability than the traditional SLAMS networks. NCore provides an opportunity to address new directions in monitoring and begin to fill measurement and technological gaps which have accumulated in the networks. Indiana operates one urban NCore site. The sites are required to measure PM2.5, speciated PM2.5, PM10c, O3, SO2, CO, NO, true NO2, NOy, and meteorology. As of June 2021, PAMS is included at NCore sites located in a CBSA with a population of 1,000,000 or more.

### Near-Road Monitoring

On February 9, 2010, the U.S. EPA promulgated monitoring regulations for the NO2 monitoring network. In the new monitoring requirements, state and local air monitoring agencies are required to install near-road NO2 monitoring stations at locations where peak hourly NO2 concentrations are expected to occur within the near-road environment in larger urban areas. Site selection is required to consider traffic volumes, fleet mix, roadway design, traffic congestion patterns, local terrain, and meteorology in determining placement of a required near-road NO2 monitor. Indiana operates one near-road monitoring site. IDEM worked with the Indiana Department of Transportation (INDOT) to obtain a location for the site. The near-road site is required to measure NO2, CO, O3, and meteorology. Toxics volatile organic compounds (VOCs), PM1.0, PM2.5, and continuous speciation (black carbon/UVC) are also measured at this site.

# Chemical Speciation Network (CSN)

As part of the PM2.5 NAAQS review completed in 1997, the U.S. EPA established a PM2.5 CSN consisting of special trends network (STN) sites and supplemental speciation sites. The CSN is a component of the national PM2.5 monitoring network. The goal of the national PM2.5 monitoring network is to monitor whether the NAAQS are being attained. However, CSN data are not used for attainment nor nonattainment decisions but intended to complement the activities of the larger gravimetric PM2.5 measurement network component. CSN data are used for multiple objectives, including:

- Assessing data trends.

- Developing effective SIPs and determination of regulatory compliance.

- Developing emission control strategies and tracking progress of control programs.

- Aiding in the interpretation of health studies by linking effects to PM2.5 constituents.

- Characterizing annual and seasonal spatial variation of aerosols.

- Comparing with chemical speciation data collected from Interagency Monitoring of Protected Visual Environments Network (IMPROVE network).

## Section 7:	Quality Objectives and Criteria for Measurement Data

## The primary data quality objective is to ensure that the data collected by the AMB are consistent, of known and adequate quality, supported by adequate calibrations and evaluations, and sufficiently complete to describe the atmospheric state with respect to spatial and temporal distribution. Minimum QA requirements are listed in 40 CFR Part 58 and its appendices. The data must meet the quality goals for representativeness, precision, bias, detectability, completeness, and comparability. Accuracy has been a term frequently used to represent closeness to “truth” and includes a combination of precision and bias error components. The term had been used throughout the CFR but has since been replaced with bias when the ability to distinguish precision from bias exists. The quality system for the AMB air monitoring program focuses on understanding and controlling, as much as possible, measurement uncertainty and by mainly focusing on the measurement quality goals for better measurements.  Representativeness is addressed through network designs and is not something that the quality system can control through better measurements.

## Collecting quality data begins with properly trained staff and adequate funding to provide the necessary equipment which meets required performance specifications. High quality data also relies on having adequate supplies available, safe monitoring locations which meet U.S. EPA siting requirements, and up to date QAPPs and SOPs.

## Table 9 lists the data quality indicators (DQI) to achieve the data quality objectives and how the AMB approaches each one.

Table 9. Data Quality Indicator/Approach

| DQI | Approach |
| --- | --- |
| Representativeness | Representativeness is the degree to which data accurately and precisely represent a characteristic of a population, a parameter variation at a sampling point, a process condition, or an environmental condition. The data collected will represent ambient air to which the public is exposed. Monitoring locations are selected to meet this objective and adhere to U.S. EPA requirements for siting. All particulate inlets must be the proper height above the ground, and have a minimum distance from objects which could affect the representativeness of the data collected as described in 40 CFR Part 58 Appendix E. Examples of siting are found in the QAS SOP Site Evaluation Procedures and include: - Minimum of 10 meters from the dripline of trees - Distance from sampler to any obstruction must be twice the height that the obstruction protrudes above the sampler - Inlet height 2 to 15 meters above ground except PM10 microscale, which is 2 to 7 meters above ground - Ground is paved or has vegetative cover Special purpose monitoring, near-roadway monitoring, and industrial-based monitoring require different siting designed to meet these special monitoring objectives. |
| Precision | Precision is a measurement of mutual agreement among individual measurements of the same property usually under prescribed similar conditions, expressed in terms of the coefficient of variation. Precision for PM2.5, PM10, and TSP (for metals) samplers is estimated using concentration measurements from collocated samplers running concurrently. Acceptable limits are provided in 40 CFR58 Appendix A. |
| Bias | Bias is the systematic or persistent distortion of a measurement process which causes errors in one direction. The PM2.5, PM10, and Pb bias estimate is based upon the monthly flow rate verifications. For Pb analysis, relative percent differences between the audit concentrations (in μg Pb per strip) and the corresponding measured concentrations (in μg Pb per strip) are used to calculate analytical bias. Acceptable limits are provided in 40 CFR58 Appendix A. |

| DQI | Approach |
| --- | --- |
| Detectability | Detectability is based on determination of the low range critical value of a characteristic which a method specific procedure can reliably discern. Maximum limits also apply. For the AMB particulate program this applies to PM2.5 and PM10 balances, metals analysis, and PM2.5 and PM10 continuous data. |
| Completeness | Data completeness is defined as the number of valid measurements (meeting all QC and QA criteria) divided by the number of possible or scheduled measurements. The AMB strives to obtain the highest level of data capture or completeness as possible. All data must meet a minimum of 75% completeness. |
| Comparability | Comparability is the degree to which different data agree or can be represented as similar. Comparability describes the confidence that two data sets can contribute to a common analysis and interpolation. Ambient air monitoring is conducted in adherence to the established methods as published in 40 CFR Part 50 and 53 for national consistency and comparability. Participation in the Performance Evaluation Program (PEP), metals round-robin, PM2.5 round-robin as well as conference calls help ensure comparability. In addition, PQAOs which submit air monitoring data into AQS will be subject to an annual third-party evaluation by the QAS staff. |

7.1	Measurement Quality Objectives

## To ensure the quality of the data, AMS staff perform sampler calibrations and verifications. These checks help determine the validity of the data by ensuring that the samplers meet specific limits. Additional performance evaluation (PE) checks are performed independently by the QAS staff. The QA results assist with calculating statistical analysis of the data and can solidify the site operator’s decision on the condition of the data. Tables 10 through 13 contain the types of checks (assessment method) and limits (measurement quality objectives) for each type of particulate sampler. The procedures for this work are provided in the AMB SOPs listed in section 6.2 of this QAPP.

Table 10. AMS Intermittent Methods Data Assessment Requirements

| Type of Check: Calibration and Verification |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measurement Quality Objectives |
| PM2.5 | Check of sampler flow rate, AT, FT, BP, external leak check, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT and FT <±2.1ºC; BP <±10.1 mmHg; external leak check <25 mmHg; time <5 minutes of local standard time; date is correct. NOTE: Flow must be calibrated <±2.1% of standard flow |

| Type of Check: Calibration and Verification |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measurement Quality Objectives |
| PM10 and PM10c | Check of sampler flow rate, AT, FT, BP, external leak check, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT/FT <±2.1ºC; BP <±10.1 mmHg; external leak check <25 mmHg; time <5 minutes of local standard time; date is correct. NOTE: Flow must be calibrated <±2.1% of standard flow |
| TSP for metals | Check of sampler flow rate, AT, BP, time, date | Flow <±7.1% of transfer standard flow and must be in the range of 1.1 to 1.7 m3/min. actual conditions; AT <±2.1ºC; BP <±10.1 mmHg; time <5 minutes of local standard time; date is correct; if using a Thermo Scientific 2025 for TSP, follow the PM2.5 Measured Quality Objectives. NOTE: Flow must be calibrated <±5.1% of standard flow |
| PM2.5 Speciation (SASS/Super SASS/URG) | Check of sampler flow rate, AT, FT, BP, leak check, time, date | Flow <±10.1% of transfer standard flow and must be in the range of 6.03 to 7.37 LPM actual conditions for SASS/Super SASS and 19.8 to 24.2 LPM actual conditions for URG; AT and FT <±2.1ºC; BP <±10.1 mmHg; leak check ≤0.1 LPM (SASS or Super SASS) and <225 mmHg (URG); time <5 minutes of local standard time; date is correct NOTE: Flow must be calibrated if verification is not <±5.1% |

Table 11. QAS Intermittent Methods Data Assessment Requirements

| Type of Check: Audit |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measurement Quality Objectives |
| PM2.5 | Check of sampler flow rate, AT, FT, BP, external leak check, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT and FT <±2.1ºC; BP <±10.1 mmHg; external leak check <25 mmHg; time <5 minutes of local standard time; date is correct |
| PM10 and PM10c | Check of sampler flow rate, AT, FT, BP, external leak check, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT and FT <±2.1ºC; BP <±10.1 mmHg; external leak check <25 mmHg; time <5 minutes of local standard time; date is correct |

| Type of Check: Audit |  |  |  |
| --- | --- | --- | --- |
| Parameter Method | Assessment Method |  | Measurement Quality Objectives |
| TSP for metals |  | Check of sampler flow rate, AT, BP, time, date | Flow <±7.1% of transfer standard flow and must be in the range of 1.1 to 1.7 m3/min. actual conditions; AT <±2.1ºC; BP <±10.1 mmHg; time <5 minutes of local standard time; date is correct; if using a 2025 for TSP, follow the PM2.5 Measured Quality Objectives. |
| PM2.5 Speciation (SASS, Super SASS, URG) |  | Check of sampler flow rate, AT, FT, BP, leak check, time, date | Flow <±10.1% of transfer standard flow and must be in the range of 6.03 to 7.37 LPM actual conditions for SASS and Super SASS and 19.8 to 24.2 LPM actual conditions for URG; AT and FT <±2.1ºC; BP <±10.1 mmHg; leak check ≤0.1 LPM (SASS and Super SASS) and <225 mmHg (URG); time <5 minutes of local standard time; date is correct NOTE: A verification is required if audit flow is not <±5.1% |

Table 12. AMS Continuous Methods Data Assessment Requirements

| Type of Check: Calibration and Verification |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measurement Quality Objectives |
| PM1.0 | Check of sampler flow rate, time, date | Flow <±5.1% of design flow checked with transfer standard; time <5 minutes of local standard time; date is correct; if using a T640 for PM1.0, follow the continuous PM2.5 Measured Quality Objectives |
| PM2.5 | Check of sampler flow rate, AT, BP, leak check, zero, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT <±2.1ºC; BP <±10.1 mmHg; leak check <1.5 LPM for basic and <0.3 LPM for advanced (BAM 1020), ≤0.2 μg/m3 (TAPI T640 and T640x); time <5 minutes of local standard time; date is correct; NOTE: Flow must be calibrated <±2.1% of standard flow |
| PM10c | Check of sampler flow rate, AT, BP, leak check, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT <±2.1ºC; BP <±10.1 mmHg; leak check ≤0.2 μg/m3 (T640x); time <5 minutes of local standard time; date is correct. NOTE: Flow must be calibrated <±2.1% of standard flow |
| PM10 | Check of sampler flow rate, AT, BP, leak check | Flow < ±7.1% of transfer standard flow; AT <±2.1ºC; BP <±10.1 mmHg; leak check ≤0.15 LPM (main flow TEOM), ≤0.60 LPM (aux flow TEOM); time <5 minutes of local standard time; date is correct |

| Type of Check: Calibration and Verification |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measurement Quality Objectives |
| PM2.5 speciation (Aethalometer) | Check of sampler flow rate, AT, BP, leak check, time, date | Flow <±10.1% of transfer standard flow; AT <±2.1ºC; BP <±10.1 mmHg; leakage test <10.0% (TAPI 633 and AE33); <0.5 lpm (BC-1060); time <5 minutes of local standard time; date is correct |

Table 13. QAS Continuous Methods Data Assessment Requirements

| Type of Check: Audit |  |  |
| --- | --- | --- |
| Parameter Method | Assessment Method | Measurement Quality Objectives |
| PM1.0 | Check of sampler flow rate, time, date | Flow <±5.1% of design flow checked with transfer standard; time <5 minutes of local standard time; date is correct; if using a T640 for PM1.0, follow the continuous PM2.5 Measured Quality Objectives |
| PM2.5 | Check of sampler flow rate, AT, BP, leak check, zero, time, date | Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT <±2.1ºC; BP <±10.1 mmHg; leak check <1.5 LPM (BAM), ≤0.2 μg/m3 (TAPI T640 and T640x); time <5 minutes of local standard time; date is correct |
| ## PM10c | ## Check of sampler flow rate, AT, BP, leak check, time, date | ## Flow <±4.1% of transfer standard flow and <±5.1% of design flow; AT <±2.1ºC; BP <±10.1 mmHg; leak check ≤0.2 μg/m3 (T640x); time <5 minutes of local standard time; date is correct |
| ## PM10 | ## Check of sampler flow rate, AT, BP, leak check, time, date | ## Flow < ±10.1% of transfer standard flow; AT <±2.1ºC; BP <±10.1 mmHg; leak check ≤0.15 LPM (main flow TEOM), ≤0.60 LPM (aux flow TEOM); time <5 minutes of local standard time; date is correct |
| ## PM2.5 speciation (Aethalometer) | ## Check of sampler flow rate, AT, BP, leak check, time, date | ## Flow <±10.1% of standard flow; AT <±2.1ºC; BP <±10.1 mmHg; leakage test <10.0% (TAPI 633 and AE33); <0.5 lpm (BC-1060); time <5 minutes of local standard time; date is correct |

## 7.1.1	Field Blanks

Field blanks for intermittent PM2.5 and PM10 are collected at a frequency of 10% of the sampling runs scheduled per site. If the field blank was properly programmed and ran as a field blank, then it will count towards the 10% collection frequency. The validation acceptance criterion for field blanks is ±30 µg between weighings. If the limit is exceeded, further investigation is performed, which may result in additional field blanks being deployed and potentially a QA qualifier or a null data qualifier being placed on any samples in question. Further information on field blanks is provided in 40 CFR Part 50 Appendix L and QA GD 2.12. Field blanks for intermittent PM2.5 speciation follow the U.S. EPA sampling schedule and follow U.S. EPA SOPs. Field blanks for TSP follow the PEP audit program.

## 7.2	PEP Audits

## PM2.5 and Pb sampler bias are determined from the results of independent Performance Evaluation Program (PEP) audits. The U.S. EPA or designated contractor performs the audits. The audits are required as part of 40 CFR Part 58 Appendix A. Requirements for the PEP include:

- For each PQAO with five or less PM2.5 and Pb monitoring sites, five valid audits per year distributed across four quarters.

- For each PQAO with greater than five PM2.5 and five Pb sites, eight valid audits per year distributed across the four quarters.

- One hundred percent completeness.

- All samplers subject to an audit within six years.

## Results of audit samples are reported to AQS by the U.S. EPA or the designated contractor.

## 7.3	Sampler Collocation

## 7.3.1	PM2.5 or PM10

Precision is the measure of mutual agreement among individual measurements of the same property, usually under prescribed similar conditions. Precision is estimated by using a duplicate sampler of the same method (e.g., 2025 to 2025; or T640 to T640) or a collocated intermittent sampler at a selected continuous monitoring location (e.g., 2025 to T640) in a measurement network.  One sampler is designated as the reporting sampler and one sampler is designated as the collocated sampler. In the case of the duplicate sampling method, the collocated sampler must be maintained, operated, calibrated, verified, and audited in the same time frame and in the same manner as the reporting sampler.

The collocated sampler’s inlet must be within 1 to 4 meters from the inlet of the reporting sampler’s inlet, within one meter vertically, and must be at least two meters from the inlet of any other sampler inlets such as high-volume samplers (e.g., TSP). IDEM’s ANP provides the number of PM2.5 (intermittent and continuous) duplicate sites, PM10 (intermittent) duplicate sites, and PM2.5 (intermittent to continuous method) sites with collocated samplers, which meets the numbers required by the CFR and U.S. EPA. In addition, collocated intermittent PM samplers are used at sites to test the comparability of the continuous PM samplers either newly installed at existing sites or at newly established sites.

Precision is calculated from the difference in the concentrations from the reporting and collocated samplers over a calendar quarter. All collocated intermittent samplers operate on a 1-in-6 day frequency except for PM2.5 continuous collocated samplers, which operate continuously every day. For each intermittent sampler, this allows for approximately 15 data pairs (reporting & collocated concentrations) over each quarter for each site with collocated samplers. Estimates of network precision are made from reviewing three years of data.

Data are reported to the U.S. EPA AQS database for both the reporting and collocated sampler, regardless of concentration. However, CV is calculated only from data pairs (reporting and collocated concentrations) when both values are equal to or above 3 μg/m3. In addition, for the continuous methods, all 24 hours for a day must be valid for the daily average to be used to calculate CV.

Equations used to calculate precision from reporting and collocated data pairs are stated in 40 CFR Part 50 Appendix L and 40 CFR Part 58 Appendix A. The U.S. EPA also provides a downloadable spreadsheet, Data Assessment Statistical Calculator, to calculate the CV. The spreadsheet can be obtained through AMTIC Ambient Air Monitoring Quality Assurance web page.

The CV determines what corrective action, if any, is needed. A precision data quality indicator of 10% CV is used for comparison purposes, which is based upon the evaluation of 3 years of collocated precision data. CV values ≥10.1% may occur within the 3-year period. However, single collocated pairs with relative percent differences ≥10.1% may require a re-weigh. If a single collocated pair is causing the quarterly CV to be ≥10.1%, then further investigation may warrant application of a QA qualifier on the reporting value or a null data qualifier to one or both the reporting and collocated samples. The AMS staff and QAS staff are alerted by the ATS staff when filter weights indicate mass differences greater than 10%. Besides reweighs, operational solutions are investigated.

Besides collocated monitoring, precision checks are obtained using filter replicate weighing. During laboratory pre-weighing and post-weighing sessions, a filter from a batch is selected for a second weighing. The acceptable limit for the difference between the first post-weight and the second post-weight is 15 µg for clean filters and exposed filters. Failure may be due to transcription errors, microbalance malfunction, or the samples failing to reach equilibrium. Other QC checks (balance standards and lab blanks) will eliminate microbalance malfunction. If the replicate weighing does not meet the criterion, a second sample is selected and re-weighed as a second duplicate check. If this second check fails the acceptance criterion and the possibility of balance malfunction and transcription errors have been eliminated, all samples in the batch are equilibrated for an additional 24 hours and re-weighed. Corrective actions continue until duplicate weights for the batch meet acceptance criteria.

## 7.3.2	Metals

Precision of the TSP monitoring for metals network is measured by use of duplicate or collocated samplers. 40 CFR Part 58 Appendix A, Section 3.4, has the following requirements for a PQAO:

- Have 15 percent of the primary monitors (not counting non-source oriented NCore sites in PQAO) collocated. Values of 0.5 and greater round up.

- Have at least one collocated quality control monitor if the total number of monitors is less than three.

- The first collocated Pb site selected must be the site measuring the highest Pb concentrations in the network. If the site is impractical, alternative sites approved by the U.S. EPA Regional Administrator, may be selected. If additional collocated sites are necessary, collocated sites may be chosen to reflect average ambient air Pb concentrations in the network.

- The two collocated monitors must be within 4 meters (inlet to inlet) of each other and at least two meters apart for flow rates greater than 200 LPM or at least one meter apart for samplers having flow rates less than 200 LPM to preclude airflow interference.

- Sample the collocated quality control monitor on a 1-in-12-day schedule. Report the measurements from both primary and collocated quality control monitors at each collocated sampling site to AQS. The calculations for evaluating precision between the two collocated monitors are described in section 4.2.1 of 40 CFR Part 58 Appendix A.

IDEM adheres to these requirements and operates at least one collocated TSP sampler for Pb. The collocated sampler operates on a 1-in-6-day schedule. The collocated sampler inlet is located between two to four meters from the reporting sampler inlet, and within one meter vertically. While the reporting sampler’s data are used for reporting air quality, the collocated sampler’s data are used to calculate network precision (CV). Calibrations, verifications, audits, sampling frequency and duration, and analysis are conducted in the same manner for each sampler. The percentage differences between the Pb concentrations (µg/m3) from the two samplers are used to determine precision. However, for precision assessments, the minimum concentration level for both samplers must be ≥.02 g/m3. The calculations are described in detail in 40 CFR Part 58   Appendix A, Section 4.2.

The CV will determine what corrective action, if any, is needed. The precision data CV quality objective of 20% is based upon the evaluation of three years of collocated precision data. CV values of ≥20.1% may occur within the 3-year period. However, if a single collocated pair is causing the quarterly CV to be ≥20.1%, then further investigation may warrant application of a QA qualifier on the reporting value or a null data qualifier for both the reporting and collocated samples. The AMS staff and QAS staff are alerted by the ATS staff when concentration differences exceed 20%. Besides reanalysis, operational solutions are investigated.

## 7.4	Metals Strip Audits

Metals strip audits consist of the analysis of glass fiber strips impregnated with different masses of Pb, Mn, and As. For Pb, six audit strips, prepared by the QAS staff, are required to be analyzed by the ATS each quarter. Although Mn and As have no federal requirement, the QAS staff treats these in a similar manner as Pb. The QAS may also apply all three metal solutions to the same strip, with one strip being the low concentration range and one strip being the high concentration range. The QAS staff uses a manufacturer’s certified reference standard solution to prepare the strips. The QAS staff also uses its own manufacturer certified pipettes. The audits are comprised of three strips at each of two different concentration ranges plus a blank. Only the low and high concentration results are reported to AQS. The lower concentration range is 9-30 g/strip for Pb, 5-13 g/strip for Mn, and 10-30 g/strip for As. The higher concentration range is 60-90 g/strip for Pb, 13-18 g/strip for Mn, and 35-55 g/strip for As. The strips are provided to the ATS monthly. Six strips, one low and one high concentration for each metal, as well as one blank are analyzed each month (if putting all metals on same strip then the monthly total would be two strips plus one blank). Results are sent to the QAS staff and evaluated. If results are ≥10.1% action is taken by the ATS staff to see whether the difference can be reduced. This may require examining the ATS equipment to determine whether maintenance is needed or re-run the samples and report back the findings to the QAS staff. Data can potentially be impacted with a QA data qualifier or a null data qualifier although each investigation into the differences dictates the specific actions taken and the values reported. Two additional Pb strips, which are obtained by the QAS staff from a U.S. EPA contracted laboratory, are also provided to the ATS for analysis each month as part of the PEP program. Results of these findings are treated the same as the strips prepared by the QAS staff and reported to AQS.

Section 8:	Training

Formal staff training is scheduled for new staff and to periodically update staff’s skills and program operations. Required training consists of QAPP and SOP review, air monitoring instrument manuals and corresponding SOPs, and hands-on training on air monitoring instruments which are related to the required work for specific staff (e.g., AA/GF for ATS staff, PM2.5 FRM for AMS staff, O3 verifications for QAS staff). Formal staff training is coordinated with the section chiefs, senior level staff, or parameter specialists in the AMS, QAS, and ATS of the AMB on an as needed basis for staff engaged in the following: operating, calibrating, verifying, validating, and auditing analyzers or samplers; laboratory procedures; field duties; safety; and any other items related to work performed by staff in the AMB. Staff training is tracked and documented by the individual section chiefs. The exception, any in-house computer safety, is documented by the Indiana State Personnel Department but able to be tracked by individual section chiefs. Computer safety training can be tracked by individual section chiefs. Standard literature references are readily available to all staff to access, whether through the internet or saved on an individual’s home drive, including the Federal Register, manufacturer’s instrument manuals, and QA GDs related to the program objectives. Courses and other training are also provided through U.S. EPA and vendors (e.g., Air Knowledge).

## Section 9:	Documentation and Records

## The goal of IDEM is to collect data which are accurate and representative of the actual conditions. For this to occur, documentation and record keeping must be performed at a high level of accuracy and consistency amongst all the PQAOs. The AMB shared drive is only available to AMB staff which helps secure against any tampering issues. Documents on the IDEM InfoDump can only be seen by IDEM staff. Documents on the IDEM InfoDump and internet can only be changed by the IDEM approved web managers. Documentation in DMDS, which follows prescribed templates, is secure and cannot be changed once entered. Data in DMDS can only be changed by the AMS(s) chiefs and staff. The procedure to access data in DMDS is provided in the AMB SOP “Gaseous Data Validation Using DMDS”. The QA laboratory cabinet is in a secure location with limited access to others. Table 14 summarizes what documentation is involved, location of the documents, retention time, and the main custodian. All records are either kept at the minimum requirements as addressed in the IDEM QMP or kept indefinitely.

Table 14. Documentation and Records

| Document | Location | Retention Time | Custodian |
| --- | --- | --- | --- |
| ANP; 5-Year network plan; QAPP | IDEM internet and InfoDUMP; AMB shared drive | Latest on IDEM internet and InfoDUMP; AMB shared drive maintains previous versions | ANP and 5-Year network plan – AMS (1 and 2) chief(s); QAPP – QAS chief and OPS QA |
| SOPs | SharePoint IDEM QA Library; AMB shared drive | Latest on SharePoint IDEM QA Library; archived versions in SharePoint archived QA Documents; AMB shared drive maintains previous SOPs | QAS program coordinator and OPS QA |
| ## Logs | ## On-site computer hard drive if available; operator log available through DMDS; Intermittent PM2.5, Intermittent PM10, and intermittent PM2.5 speciation have paper logbooks kept at the site | ## Operator logs kept indefinitely; Existing site logbooks kept indefinitely; If a paper logbook runs out of space or the sampler is removed from service, the logbook is kept at AMS lab indefinitely; Discontinued site logbooks retained 7 years | ## AMS DMDS administrator for continuous data; AMS parameter specialist for intermittent data |
| ## Intermittent PM2.5 and PM10 Data/Chain of Custody Sheets | ## Binders (organized by site) on bookshelf, and then transfer to VFC | ## 1 year in binders; indefinitely in VFC | ## ATS environmental chemist |
| ## Intermittent PM2.5 and PM10 Gravimetric Data | ## GLIMS | ## Minimum of 3 years | ## ATS chief |

| ## Document | ## Location | ## Retention Time | ## Custodian |
| --- | --- | --- | --- |
| ## Intermittent PM2.5 and PM10 Calibration and Verification Completed Forms | ## File cabinets in AMB office area | ## Kept indefinitely | ## AMS parameter specialist |
| ## Intermittent PM2.5, PM10, TSP, and PM2.5 Speciation Audit Forms | ## AMB Shared Drive | ## Kept indefinitely | ## QAS parameter specialist |
| ## Intermittent PM2.5, PM10 and speciation field logbooks | ## Kept with sampler | ## Stays at the site unless book runs out of space or sampler is removed from service; then kept at AMS Lab indefinitely | ## AMS parameter specialist |
| ## Intermittent TSP data card | ## ATS laboratory for 1 year, then to warehouse | ## 5 years | ## ATS chief |
| ## Intermittent metals lab analysis data | ## AMB shared drive | ## AMB shared drive maintains previous years of data | ## ATS chief |
| ## Intermittent TSP calibration and verification forms | ## Kept with parameter specialist and on AMB shared drive | ## Kept indefinitely | ## AMS parameter specialist |
| ## Intermittent PM2.5 speciation data sheet; intermittent PM2.5 speciation calibration and verification Forms | AMS files | Minimum of 3 years | AMS environmental manager |
| ## Continuous PM1.0, PM2.5, PM10, and PM2.5 speciation Data | DMDS | DMDS maintains data indefinitely | AMS DMDS administrator for continuous data |
| ## Continuous PM1.0, PM2.5, and PM10 calibration forms | AMS files | Kept indefinitely | AMS parameter specialist |
| ## Continuous PM1.0 verification form | AMS files; AMB shared drive | Kept indefinitely | AMS parameter specialist |
| ## Continuous PM1.0, PM2.5, PM10, and PM2.5 speciation audit forms | AMB shared drive | Kept indefinitely | QAS parameter specialist |

| ## Document | Location | Retention Time | Custodian |
| --- | --- | --- | --- |
| ## Continuous PM2.5 and PM10 verification forms | AMS Files; AMB shared drive | Kept indefinitely | AMS parameter specialist |
| ## Continuous PM2.5 speciation data disk | Data put on AMB shared drive and uploaded to DMDS after validation; disk then reused | DMDS maintains data indefinitely | AMS parameter specialist |
| ## Continuous PM2.5 speciation calibration and verification forms | AMS files | Minimum of 3 years | AMS parameter specialist |
| ## Continuous PM2.5 verification Form | AMS files; AMB shared drive | Kept indefinitely | AMS parameter specialist |
| ## PM2.5 and PM10 exposed filters | Refrigerator; then warehouse | 1 year in refrigerator warehouse 5 years then disposed | ATS chief |
| ## TSP for metals exposed filters | ATS laboratory for 1 year, then to warehouse | 5 years | ATS chief |
| ## PM2.5 BAM tape | AMS laboratory | New – expiration date; exposed - 1 year | AMS parameter specialist |
| ## Quarterly comparison of weights used for PM gravimetric | On balance computer | 5 years | ATS environmental chemist |
| ## PM diary | ## Logbook in lab; pre-2011 archived on balance computer | ## 5 years | ## ATS environmental chemist |
| ## PM data memos generated by ATS | ## Send via e-mail; not saved by ATS | ## N/A | ## QAS chief |
| ## PM2.5 and PM10 balance service logs | ## Stored on the instrument PC then moved to shared drive | ## Kept indefinitely | ## ATS environmental chemist |

| ## Document | ## Location | ## Retention Time | ## Custodian |
| --- | --- | --- | --- |
| ## Atomic absorption method detection limits as per 40 CFR Appendix B to Part 136 | ## AMB shared drive | ## Kept indefinitely | ## ATS environmental chemist |
| ## Calibrations, certifications, and verifications performed by the QA laboratory | ## AMB shared drive | ## Information kept at least 3 years unless item is still in circulation then kept indefinitely | ## QA laboratory manager |
| ## NIST-traceable certifications | ## QA laboratory cabinet file | ## Indefinitely | ## QA laboratory manager |
| ## TSA questionnaire, audit finding response form, and final report and closeout letter | ## AMB shared drive | ## Kept indefinitely | ## QAS Chief and QAS Program Manager |

## Section 10:	Network Description (or Sampling Process Design)

The primary rationale for the operation of the PM network is to determine compliance with the NAAQS and to provide the public with information on current air quality. Particulate sampling is primarily conducted in population centers per U.S. EPA requirements. Additional sites are also operated to establish rural background concentrations, provide statewide as well as regional coverage, and for areas of specific interest, such as sources and near road locations. The IDEM ANP provides information on sites and can be found at https://www.in.gov/idem/airquality/.

Network design and sampler siting are established based on 40 CFR Part 58 Appendices A (collocation), D (design criteria) and E (siting criteria) and are mentioned in Sections 6 and 7 of this QAPP.

When photographs or digital images are taken for purposes of documenting and supporting a field investigation, such as a QAS site evaluation, a record of each exposure or image is saved as a file on an AMB computer shared drive. The following information is recorded:

- The site name and subject of the image will be part of the file’s name, which is stored based on the year it was taken. For example, file name “Gary IITRI SE” stored under the path QA\Site Information\Site Evaluations\Gary IITRI\2020.

- The name of the staff taking the photograph or digital image will correspond to any paperwork, such as a site evaluation form. If no paperwork is used, a log entry in DMDS is adequate.

Section 11:	Sampling Method Requirements

Sampling equipment and procedures follow 40 CFR Part 50 Appendices B, G, J, and L. Specific instructions on technical aspects of the procedures can be found in the specific sampler’s manual as well as the following AMB SOPs:

- 2025 A/B Sequential Air Sampler Calibration and Maintenance

- 2025 A/B Sequential Air Sampler Verification and Maintenance

- Met One Instruments Beta Attenuation Monitor (BAM) 1020 Calibration and Maintenance

- Met One Instruments Beta Attenuation Monitor (BAM) 1020 Verification and Maintenance

- Teledyne API 633 Aethalometer Verification and Maintenance

- TAPI T640 PM Mass Monitor Verification and Maintenance

- TAPI T640 PM Mass Monitor Calibration and Maintenance

- 2025i Sequential Air Sampler Verification and Maintenance

- 2025i Sequential Air Sampler Calibration and Maintenance

- Tisch Environmental Hi-Vol+ Lead Sampler Calibration and Maintenance

- Tisch Environmental Hi-Vol+ Lead Sampler Verification and Maintenance

- URG 3000N Calibration and Maintenance

- URG 3000N Verification and Maintenance

## The AMB maintains a complete set of SOPs for all procedures, which are available through the AMB shared computer drive and the IDEM InfoDUMP.

## 11.1	Sampling Equipment

## The AMB utilizes samplers which meet established federal reference or equivalent method requirements except for some particulate sampling designated as non-criteria (Table 15). Documentation of changing out instrumentation or a method change at a site is made in the site DMDS program.

Table 15. PM Sampling Equipment

| Parameter | Reference Method | Monitor Type | Sample Frequency |
| --- | --- | --- | --- |
| PM1.0 continuous | No federal reference method | TSI 3783 | Continuous |
| PM1.0 continuous | Automatic equivalent method: EQPM-0516-236 | TAPI T640 | Continuous |
| PM2.5 intermittent | Manual reference method: RFPS-0498-118 | Thermo Scientific 2025 A/B Thermo Scientific 2025i | Manual, 3-day, and 6-day |
| PM2.5 continuous | Automatic equivalent method: EQPM-0516-236 | TAPI T640 | Continuous |
| PM2.5 continuous | Automated equivalent method: EQPM-0516-238 | TAPI T640X | Continuous |
| PM2.5 continuous | Automated equivalent method: EQPM-0308-170 | Met One Instruments BAM 1020 | Continuous |
| PM10c intermittent | Manual reference method: RFPS-0509-176 | Thermo Scientific 2025 A/B Thermo Scientific 2025i | Manual, 3-day or 6 -day |
| PM10c continuous | Automated equivalent method: EQPM-0516-240 | TAPI T640X | Continuous |
| PM10 intermittent | Manual reference method: RFPS-1298-127 | Thermo Scientific 2025 A/B Thermo Scientific 2025i | Manual, 3-day, and 6-day |
| PM10 continuous | Automated equivalent method: EQPM-0516-239 | TAPI T640X | Continuous |
| PM10 continuous | Automated equivalent method: EQPM-1090-079 | Thermo Scientific 1405 | Continuous |
| TSP for metals intermittent | Manual reference method: 40 CFR Part 50, Appendix B | Tisch Environmental Sampler HV-Plus Thermo Scientific 2025 A/B BGI PQ100 | Manual, 6-day |
| PM2.5 speciation intermittent | No federal reference method | Met One Super SASS and SASS; URG 3000N | Manual, 3-day at PAMS, 6-day at all other sites |
| Parameter | Reference Method | Monitor Type | Sample Frequency |
| PM2.5 speciation continuous | No federal reference method | Magee Scientific AE33; Met One BC-1060; TAPI 633 Dual Channel Aethalometer | Continuous |

## 11.2	Sampling Methodology

## A summary of the methods for sampling each particulate follows. Specific concepts and procedures are provided in each instrument’s operating manual.

## 11.2.1	Intermittent

- PM2.5 – The sequential sampler draws ambient air at a flow rate of 16.67 lpm through a sample inlet (size selective for PM10) and through a secondary size-selective device (selective for PM2.5). Samples are collected on pre-weighed 46.2 mm TeflonTM filters over a 24-hour period. Data are collected from the sampling events on paper records and manually entered into an operator log. Data from the sampler are also downloaded monthly. Upon return from the field, filters are re-acclimated to clean room conditions and then re-weighed in an AMB clean room(s).

- PM10, PM10c – The sequential sampler draws ambient air at a flow rate of 16.67 lpm through a sample inlet (size selective for PM10). Samples are collected on pre-weighed 46.2 mm TeflonTM filters over a 24-hour period. Data are collected from the sampling events on paper records and manually entered as an operator log. Data from the sampler are also downloaded monthly. Filters are weighed in the AMB clean room(s) following the same protocol as for PM2.5 samples. If required, PM10c concentrations are calculated from the difference between PM10 and PM2.5 concentrations.

- Metals – Samples are collected on 8” x 10” high purity glass microfiber filters over a 24-hour period. Data are collected from the sampling events on paper records and manually entered into an operator log. Data from the sampler are also downloaded monthly. Analysis for metals is performed at the ATS’s Laboratory utilizing U.S. EPA method EQL-0895-107, flameless atomic absorption spectrometry method.

- PM2.5 Speciation – Samples are collected using the SASS, Super SASS, and the URG 3000N. The SASS and Super SASS each use two separate filter media to collect PM2.5 samples. Each medium is analyzed separately for different components; a TeflonTM filter for thirty-three (33) trace metals, using Energy Dispersive X-ray Fluorescence; and a nylon filter for sulfates, nitrates, and three (3) cations (ammonium, potassium, and sodium), using ion chromatography. The URG uses a quartz filter medium for carbon analysis using thermal optical reflectance. Data are collected from the sampling events on paper records as well as data downloads. Filters are analyzed by a U.S. EPA contractor.

## 11.2.2	Continuous

## Continuous particulate samplers (not including Met One 1060) must be in a shelter. Usually, this shelter is an IDEM air monitoring trailer. The samplers are required to operate under specific operating temperatures. Table 16 lists the temperature range required for each type of sampler to collect valid data. The standard deviation limit provides temperature stability information and is not necessarily an indication of a temperature infraction resulting in invalid data.

Table 16. Continuous Particulate Sampler Indoor Temperature Requirements

| Sampler | Temperature Range |
| --- | --- |
| Magee Aethalometer AE33 Thermo Scientific TEOM 1405 Met One Instruments BAM-1020 TAPI T640 TAPI T640X | 0.0 to 40.0 °C; < 2.1 °C SD over 24 hours |
| TSI 3783 TAPI Aethalometer 633 | 10.0 to 40.0 °C; < 2.1 °C SD over 24 hours |

- PM1.0 – The TSI 3783 monitor is a continuous laminar flow condensation particle counter which uses water as the working fluid. The monitor provides rapid, high-precision measurement of the numbers of ultrafine airborne particles, down to 7 nm. The monitor draws in an air sample through a sharp cut cyclone (SCC) and counts the number of particles in the sample to provide a particle concentration value which is displayed as the number of particles detected per cubic centimeter of sampled air. The T640 also measures PM1.0. This sampler draws air through an inlet and uses scattered light spectrometry to separate and bin particles down to 180 nm.

- PM2.5, PM10, PM10c – Ambient air is drawn into sampler through an air inlet followed by an exchangeable filter cartridge or filter tape, on which the particulate mass collects. The inlet system is equipped with a sampling head which separates particles of either a 2.5 µm or      10 µm diameter (PM2.5 or PM10). The sampled air proceeds through the sensor unit; microbalance for the TEOM, beta attenuation for the BAM, and scattered light spectrometry for the TAPI T640 and TAPI 640X. Data are collected continuously. The 640X calculates PMc.

- PM2.5 speciation – An Aethalometer is used to continuously collect a variety of carbon species. The models used in Indiana’s air monitoring network allow measurement at two wavelengths of light: 880 nm (infrared) and 370 nm (ultraviolet). This distinction aids in determining the source of carbon; diesel exhaust, wood smoke, etc. The Aethalometer uses filtration and optical measurement to provide continuous concentration data. The sample air is pulled through a quartz fiber filter tape, where the aerosol black carbon is deposited. The Aethalometer continually analyzes the sample gas while collecting the sample. Once the filter tape has reached a set attenuation, it advances. The Aethalometer passes a beam of light through the sample depositing on the filter tape. The detector measures the amount of light transmitted through the sample. This quantity is linearly proportional to the black carbon in the filter deposit.

## 11.3	Failed Sample Events

## In the event of a failed sample which was not collected according to specific requirements, for intermittent samplers the AMS staff will record information on the chain of custody form, giving detailed information of why the sample is invalid or whether a qualifier is needed. The ATS staff will investigate and apply proper QA qualifier or null code in GLMIS, and send a report to AMS, QAS, and ATS staff. If a QAS auditor finds an issue, then a memo is sent to the AMS staff, who then confirms the results. A detailed log is also entered in DMDS for the specific site, whether for an intermittent or a continuous particulate sampler. The U.S. EPA contract lab for PM2.5 intermittent speciation may also apply null codes and qualifiers, which are reviewed by the AMS staff and QAS staff.

## Section 12:	Sample Handling and Custody

12.1	Sample Handling

The actual handling of samples includes intermittent PM2.5, intermittent PM10, intermittent TSP for metals, and intermittent PM2.5 speciation. Samples are handled as specified in AMB SOPs listed in this QAPP. The major components of sample handling include labeling, sample collection, and transportation. The continuous particulate sampling provides data through telemetry as well as on data disks.

12.2	Sample Labeling

The labeling or proper marking of filters (or filter holders) and monitoring devices ensures positive identification throughout the sampling and analysis process. When ink is used to mark paperwork, sample inlets or size-selective devices, use indelible which is unaffected by the gases and temperature to which it is subjected. When errors occur on paperwork, corrections are made by drawing a single line through the error and entering the correct information. The staff making the correction must initial and date the change. If possible, all corrections are made by the staff making the error. Bar code identification is also used on filter paperwork and filter holders. An example of this is in the AMB SOP, “Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software”.

All transport containers have a unique identification which precludes the possibility of interchange. The ID number of the filter is recorded on the filter paperwork or the sample card and accompanies the sample. Properly handle the sample to ensure no contamination occurs and the sample analyzed is indeed the sample taken under the conditions reported. Any security measures taken are documented by written record.

All intermittent filters (or filter holders) are accompanied by a chain of custody form. Chain of custody forms are labeled and identified with the following information using an indelible marking device:

- Sample location

- Sample number

- Parameter(s) sampled (e.g., PM2.5, PM10, Pb)

- Time of sampling and removal

- Date of sampling and removal

- Operator signature

- Sampling equipment identification (e.g., motor number)

- Flow controlling device identification

- Sampling conditions (e.g., meteorological data, flow readings, etc.)

12.3	Sample Collection

Once the sampling process has occurred, the sample is carefully removed from the monitoring device and placed in a sealed nonreactive container. The sealing process depends on the type of container used. Sealing is done to protect the sample from accidentally being exposed, and where a possibility of tampering may exist, self-adhesive stickers are needed which can be signed by the sample handlers. Damage to filters or filter holders are noted on the chain of custody form.

12.4	Sample Transportation

During the transportation of samples and other monitoring data (e.g., SD card), eliminating the possibility of tampering, accidental destruction, and any physical or chemical action on the sample is important. Staff having the custody of the samples or other data must be able to testify that no tampering has occurred, and that security was continuously maintained. Samples are shipped in a cooler with ice packs and a certified min/max thermometer to ensure temperature limits are met.

12.5	Sample Custody

A sample or data are considered under a staff member's custody if:

- In a staff member's physical possession

- In view of the staff member after taking possession

- Secured by the staff so that no one can tamper with the sample

- Secured by the staff in an area restricted to authorized staff

The fewer people handling samples and data the better. Anyone handling such samples or data must be associated with the project. When shipping data to its reduction point, follow transportation precautions to avoid tampering.

# A logbook is kept at the site for intermittent PM2.5, PM10, and PM2.5 speciation samplers. Any work performed is documented in the logbook. All sites also have an operator log, which is kept in DMDS.

For samples on which the ATS does the analysis, laboratory data books serve as chain of custody since the following information is contained in the laboratory data books:

- Date samples were received in the laboratory

- Signature or initials of sample custodian

- Identification of samples (filter numbers and samples dates)

- Date of sample analysis and signature or initials of analyst

- Date the filter was quality assured; includes signature or initials of QAS staff member

- Parameters analyzed and results

- Disposition of samples

After completing the analysis of the PM2.5 and PM10 samples, store the samples in the original filter petri slide for 1 year in a refrigerator located in a limited access area. After one year, filters are stored in a warehouse and disposed of after 5 years. TSP filters for metals are also stored in the original filter card for 1 year in a limited access area then stored in a warehouse. After 5 years the filters are disposed.

When samples collected at one location are shipped or hand carried to another agency for analysis, chain of custody procedures are followed. All samples are accompanied by a chain of custody form. The form includes who the staff relinquishing the samples and the signature of the staff receiving the samples. All samples are hand carried or sent by a reputable courier service such as the U.S. Postal Service. Shipping packages are sealed. Once the samples are delivered to the laboratory, the addressee or a designated substitute ensure no package tampering has occurred. The addressee then opens the package and verifies the contents. The staff sign the package form that samples were or were not received in the original package and that all appropriate information has been addressed. The samples are logged in at the laboratory facility and placed in a limited access area until and during analysis. When dealing with multiple parameters in the same package container, each staff member handling the samples or portions of the sample indicates handling by signing the chain of custody form.

12.6	Continuous Monitoring Data

Unique chain of custody problems may occur with fully automated data acquisition systems. No one has possession of the data while traveling from the sampling site to the data processing facility. Reporting organizations must maintain Chain of Custody procedures for printouts. When continuous data arrive at the point of reduction (e.g., laboratory, AMB office) and a physical exchange of possession occurs, a chain of custody form must accompany the data. Forms are electronically copied into IDEM virtual file cabinet.

12.7	Limited Access

The AMS maintains all monitoring devices in an area of limited access (i.e., lock rooms or stations) and only allows access to staff whose work requires entry into the area.

## 12.8	Installation and Removal of Data Storage Cards

Anytime a continuous monitoring data storage card is installed in or removed from an analyzer, e.g., Aethalometer, the following information is documented on the card:

- Sample location

- Parameter(s) sampled

- Time of installation and removal

- Date of installation and removal

- Operator identification

## 12.9	Storage of Raw Data

After the data are reduced and quality assured, store the data for a minimum of 3 years, according to the OAQ guidelines. These guidelines are found in the Indiana Archives and Records Administration, Retention Schedules.

12.10 Make-up Samples

If for any reason a sample is invalid, whether being an operation or laboratory issue, a make-up sample is attempted before the next scheduled sample or exactly one week later. If available, any unused filters or field blank filters associated with that sampler are used before new weighed filters are sent out to a site. To capture these make-up samples, sampling frequency is changed to everyday from the one-in-three-day normal frequency. This frequency is changed back once the make-up samples are collected.  Staff will limit the number of make-up samples per sampler to 5 valid samples per quarter.

## Section 13:	Analytical Methods

The AMS staff collect filters for the determination of intermittent PM2.5, PM10, PM2.5 speciation, and TSP for metals concentrations. The filters are taken to the laboratory and prepared for analysis by the ATS. Intermittent PM2.5 speciation filters as well as chromium and other specific metals are analyzed by a U.S. EPA contracted laboratory.

13.1	Laboratory Requirements for PM2.5 and PM10

Laboratory requirements for PM2.5 and PM10 sampling includes two clean rooms for the conditioning and weighing of the intermittent PM2.5 and PM10 filters. The clean rooms are a restricted access area which meets the criteria in CFR Part 50 Appendix L and U.S. EPA’s QA GD 2.12, “Monitoring PM2.5 in Ambient Air Using Designated Reference or Class I Equivalent Methods”. The criteria include:

- Mean temperature between 20.0-23.0 °C

- Temperature controlled within ±2.0 °C standard deviation over 24 hours

- Mean humidity maintained at 30.0-40.0%

- Humidity control ±5.0% standard deviation relative humidity over 24 hours

- Temperature and relative humidity are continuously monitored with 5-minute averages, collected by a data logger, and information saved to a secure computer hard drive

- Microbalance that can weigh to the nearest 0.001 mg (that is, 1 μg or one-millionth gram)

- Anti-static devices – neutralizes static charge on filters and on chairs

- Working standard weights – NIST certified standard that is used for calibrations and routine quality control checks. At some frequency, its accuracy is checked against a primary standard

- Primary standard weights - The standard within the lab that is traceable to NIST and has “authority” over other standards (working standards) in the laboratory; meaning, the working standards are checked against the primary standard at some interval or when there is concern about the accuracy of the working standard.

In addition, NIST-traceable QAS certified relative humidity and temperature measurement instruments are maintained to audit laboratory conditions.

13.2	Filter Preparation and Analysis for PM2.5 and PM10

Upon delivery of an annual order of approved 46.2 mm TeflonTM PTFE (polytetrafluorethylene) filters, receipt of the filters is documented, then the filters are left in the original shipping container and stored in the clean room until they are needed. Storing filters in the clean room makes maximizing the amount of time available for conditioning easier. As space becomes available in the clean room laboratory cabinet, the filters are removed from the shipping container and placed in a numerical order on a shelf inside cabinet for storage. When needed, we open cases of filters in a numerical order one at a time which are already labeled with a lot number and filter numbers and placed them in the pre-conditioning section of the clean room for at least 24 hours. Filters are retrieved from the case for candling (inspecting the filter under light) to check for any damage to the filter. Damaged filters (e.g., pinholes, separation, chaff, loose material, discoloration, or filter nonuniformity) are thrown away. All acceptable filters are placed in numerical order inside individual petri slides in a tray, kept in desiccator, and allowed to equilibrate to the clean room conditions for at least 24 hours prior to labeling. Then filters are ready for sampling procedures in accordance with 40 CFR Part 50 Appendix L.

13.3	Sample Weighing for PM2.5 and PM10

The ATS uses a Mettler Model XP2U microbalance for the gravimetric analysis of the samples.  This instrument meets all criteria set forth in 40 CFR Part 50 Appendix L and is maintained under agreement with the manufacturer. Two sets of Troemner weights are used as primary and working standards. The detailed procedure for the handling and analysis of samples is found in the AMB SOP, Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software.

Both pre- and post- sample filters are weighed on the same Mettler XP2U balance. However, if one clean room is not in operation, then weighing on the other clean room’s balance may occur. All weighing is done by the same ATS laboratory staff whenever possible. Refer to Table 17 for additional information regarding filter preparation and analysis checks.

Table 17. Filter Preparation and Analysis Checks

| Activity | Method and Frequency | Requirements |
| --- | --- | --- |
| Microbalance use | Calibrate or certify annually. | Resolution of 1 µg, repeatability of 1 µg. |
| Control of balance environment | Perform QC checks to keep conditions within 2.0 ˚C SD and 5.0% RH SD. | Climate-controlled, draft-free room or equivalent. Temperature and relative humidity sensors checked quarterly by QA staff to ensure accuracy. |
| Use of mass reference standards | Calibrate or certify annually. Also check working standards quarterly against laboratory primary standards. | Standards bracket weight of filter, individual standard's tolerance is less than 25 µg, handle with smooth, clean, nonmetallic forceps. |

| Activity | Method and Frequency | Requirements |
| --- | --- | --- |
| Filter handling | Observe handling procedure. | Use smooth, clean forceps. Replace 210Po antistatic strips every 6 months. |
| Filter integrity check | Visually inspect each filter. | No pinholes, separation, chaff, loose material, discoloration, or filter non-uniformity. |
| Filter identification | Label filter handling container (petri slide) with printed barcode assigned to the individual site. Write the site name on protective container (magazine). Write the filter number (last 3 digits) on laboratory data form in permanent ink. | Make sure numbers are written legibly. |
| Pre-sampling filter equilibration | Determine the correct equilibration conditions and period, at least 24 hours, for each new lot of filters. Observe and record the equilibration chamber relative humidity and temperature. Enter on lab data form. | Check for stability of laboratory blank filter weights. Weight changes must be <15 µg before and after equilibration. Mean RH between 30.0 and 40.0%, with a variability of not more than 5.0% SD over 24 hours. Mean temperature is held between 20.0 and 23.0 °C, with a variability of not more than 2.0 °C SD over 24 hours. |
| Initial filter weighing | Observe all weighing procedures. Perform all QC checks. | Neutralize electrostatic charge on filters. Wait long enough so that the balance indicates a stable reading, oscillates no more than ±2 µg and drifts no more than 3 µg, in 5-10 sec. |
| Internal QC | Zero the microbalance after each filter and reweigh one of the working standards after every ten filters. Weigh a laboratory blank filter with each weighing. Reweigh at least one filter per 25 filters (repeat weighing). | The working standard measurements must agree to within 3 µg of the certified values. The blank and repeat measurements must agree to within 15 µg. |
| Post-sampling inspection, documentation, and verification | Examine the filter and field data sheet for correct and complete entries. If sample was shipped in a cooled container, verify that low temperature was maintained. | No damage to filter. Field data sheet complete. AMS staff followed SOP. Sampler performed appropriately. |
| Filter archival | Pack and keep in cold storage for 1 year from the sampling date. Once removed, keep onsite for a minimum of 5 years. | Cold storage (refrigerate, not freeze) for a minimum of 12 months. Then onsite storage in a clean, dry area, protected from light, vibrations, and dust sources. |

13.4	Environmental Control Requirements for PM2.5 and PM10

The temperature and relative humidity requirements for PM2.5 are explicitly detailed in 40 CFR Part 50 Appendix L, QA GD 2.12, this QAPP, and the AMB SOP “Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software”. PM10 samples, which are operated on low volume samplers, are handled with the same standards. In the clean room, the filters are conditioned for a minimum of 24 hours prior to pre-weighing. Allow exposed filters to condition a minimum of 24 hours (required) but not more than 72 hours (recommended). The clean room temperature and relative humidity requirements are listed in Table 18. During transport from the clean room to the sample location, no specific requirements for temperature control are expressed. However, the filters are in their protective container and excessive heat or cold is avoided. The temperature of the filter cassette during sampling operation and in the period from the end of sampling to the time of sample recovery shall not exceed that of the ambient temperature by more than 5 °C for more than 30 minutes. Table 18 contains the temperature requirements.

Table 18. Temperature and Humidity Requirements

| Item | Temperature and Humidity Requirements |
| --- | --- |
| Clean room | 20.0-23.0 °C; 30.0-40.0 %RH |
| Pre-weighed filter | 20.0-23.0 °C and within 2.0 °C SD, 30.0-40.0 %RH and within 5.0 %RH SD, for 24 hours prior to weighing |
| Filter temperature control during sampling and until recovery | No more than 5.0 oC above ambient temperature |
| Post sample transport so that final weight may be determined up to 10 days after end of sample period | 25.0 °C or less |
| Post sample transport so that final weight may be determined up to 30 days after end of sample period | 4.0 °C or less |

13.5	Permissible Holding Times for PM2.5 and PM10

IDEM adheres to the permissible holding times, for samples, which are clearly detailed in 40 CFR Part 50 Appendix L, QA GD 2.12, detailed in this QAPP, and AMB SOP’s. For the post sample filter storage temperature, IDEM maintains the temperature 10 days even though both time frames are listed. All holding times are listed below in Table 19.

Table 19. Holding Time Restrictions for Filters

| Item | Holding Time | Time Period |
| --- | --- | --- |
| Pre-weighed filter | ≤30 days | From date of pre-weigh to date of sample |
| Recovery of filter | ≤177 hours | From completion of sample period to time of sample recovery |
| Transport of filter | <24 Hours (ideally) | From time of recovery to time placed in conditioning room |

| Item | Holding Time | Time Period |
| --- | --- | --- |
| Post-sample filter continuously stored at ≤25.0 °C | ≤10 days once sampling stops | From sample pickup date and time to date of post weigh |
| Post-sample filter stored at ≤4.0 °C | ≤30 days once sampling stops | From sample pickup date and time to date of post weigh |

Sample recovery must be performed within 177 hours from the end of the sample period. Table 20 illustrates set-up, run, and recovery dates based on sample frequency requirements of a 1-in-3 day sampling frequency.

Table 20. Sample Set-up, run, and Recovery Example Schedule*

| Sample Frequency | Sampler Type | Sunday | Monday | Tuesday | Wednesday | Thursday | Friday | Saturday |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 in 3 Week 1 | Multiple day | Sample Day 1 |  |  | Sample Day 2 | Recovery and Set-up |  | Sample Day 3 |
| 1 in 3 Week 2 | Multiple day |  |  | Sample Day 4 | Recovery and Set-up |  | Sample Day 5 |  |
| 1 in 3 Week 3 | Multiple day |  | Sample Day 6 | Recovery and Set-up |  | Sample Day 7 |  |  |
| 1 in 3 Week 4 | Multiple day | Sample Day 8 | Recovery and Set-up |  | Sample Day 9 | Recovery and Set-up |  | Sample Day 10 |
| 1 in 3 Week 5 | Multiple day |  |  | Sample Day 11 | Recovery and Set-up |  | Sample Day 12 |  |
| 1 in 3 Week 6 | Multiple day |  | Sample Day 13 | Recovery and Set-up |  | Sample Day 14 |  |  |

*As found in the U.S. EPA SLAMS Quality Assurance Guidance Document EPA-454/R-98-005 Table 11-1 Monitoring

## 13.6	Laboratory Requirements for TSP Filters for Metals

Laboratory requirements for TSP sampling for metals include an area in the ATS laboratory where unexposed filters are inspected, and exposed filters are analyzed.

## 13.6.1	Laboratory Area

The ATS staff prepares and analyzes filters in their laboratory. Lead, Manganese, and Arsenic analysis is conducted in accordance with U.S. EPA method EQL-0895-107. Certified reference standard solutions used for the analysis are stored in the ATS laboratory. Outlined in 40 CFR Part 50 Appendices L and M are the acceptance criteria for filters used in TSP monitoring for metals. Glass fiber filters are brittle and easily damaged. Take care at all times to avoid damage or contamination. An identification number is printed on each filter by the manufacturer. If filters are mailed to a laboratory for analysis, field staff are equipped with reinforced envelopes and manila folders for the protection of the exposed filters.

All filters are inspected for defects before use. An effective method to inspect filters for defects is with a light box. A light box uses a low wattage light bulb to illuminate a clear or translucent surface. Laying an 8” x 10” filter on this lighted surface aides in the inspection process. Defective filters are not used for sampling. Some common defects include pinholes, loose material, discoloration, and non-uniform appearance.

Once unexposed filters have been inspected, place each filter back in the storage box in which they were received. Then store in the ATS laboratory until the AMS staff obtains filters for use at the site. Store exposed filters, which will not be analyzed immediately, within a protective covering to prevent damage and the loss of particulate matter. A filter holder card provides adequate protection and a means on which to record important sampling information.

Use care when removing the filter from the sampler. The glass fiber filters are brittle and easily damaged. Once the filter is removed from the sampler, it may be folded along its long axis with the exposed side facing inward and placed in the filter holder card. The AMS staff normally uses a filter cassette to transport filters to and from the site. If any pieces of the filter break loose, place the pieces in the holder card. The holder card containing the folded filter can now be placed in an envelope for transport to the laboratory for analysis. If the filter is transported in the cassette, it is not folded, and just shipped with the cover over the cassette. A cover on the cassette protects the filter from damage and loss of sample.

After returning from the site, the exposed filter is returned, while inside the card or cassette, to the AMS environmental manager in charge of metals. The AMS environmental manager then removes the exposed filter from the cassette and places the filter in the filter holder card.

Section 14:	Quality Control Requirements

Quality control and acceptance criteria including time frames, for field sampling and laboratory, are detailed in Sections 6, 7, and 13 of this QAPP. Table 21 lists the action taken for field items when results do not meet measured quality objectives. Table 22 lists the action taken for laboratory items when results do not meet measured quality objectives. SOPs mentioned throughout this QAPP describe the procedures. The SOPs are available on the AMB shared computer drive and in the SharePoint IDEM QA Library site.

Table 21. Measured Quality Objective Checks and Outcomes

| Check | Outcome |
| --- | --- |
| AMS calibration frequency | If more than 13 months have passed (6 months for TSP) since the last calibration date, data will be assigned a QA qualifier “1” if verifications and other checks indicate accurate data. If additional checks are not available, data are invalid from 13 months (6 months for TSP) after the calibration until a new calibration is performed. Data are replaced with an “EC” null data qualifier. |

| Check | Outcome |
| --- | --- |
| AMS verification frequency | If a verification is missed, data may be assigned a QA qualifier “1” depending on further checks of the data, site inspections, QAS audits, etc. Any indication of an issue may invalidate data starting from the last passing verification up to a new passing verification. |
| AMS PM1.0, PM2.5, PM10, TSP, PM2.5 speciation flow verification | If a verification fails, data are invalidated using an appropriate null data qualifier back to the last passing flow verification unless sufficient evidence exists to indicate otherwise. |
| AMS PM2.5, PM10, TSP, PM2.5 speciation AT verification | For PM2.5, PM10, and URG speciation, data are invalid back to the last passing AT verification using appropriate null data qualifier, unless sufficient evidence exists to indicate otherwise. For TSP, SASS, and Super SASS speciation, data suspect back to last passing AT verification. Further evidence may require a QA qualifier or a null data qualifier. |
| AMS PM2.5, PM10, PM2.5 speciation FT verification | For PM2.5 and PM10, data are suspect back to the last passing FT verification. Further evidence may require a QA qualifier or a null data qualifier. For SASS and Super SASS speciation, data invalid back to the last passing FT verification using appropriate null data qualifier, unless sufficient evidence exists to indicate otherwise. |
| AMS PM2.5, PM10, TSP, PM2.5 speciation BP verification | Data are suspect back to last passing BP verification. Further evidence may require a QA qualifier or a null data qualifier. |
| AMS PM2.5, PM10 external leak check | Data are suspect back to last passing external leak check. An internal leak check must be performed. |
| AMS PM2.5, PM10 internal leak check | Data are invalid back to last passing internal or external leak check using appropriate null data qualifier. |
| AMS PM2.5 span dust check | AMS will adjust the peak channel number of the sampler. |

| Check | Outcome |
| --- | --- |
| AMS PM2.5 speciation leakage test | Data are screened back to last passing leakage test. Further evidence may require a QA qualifier or a null data qualifier. |
| QAS audit frequency | If an audit is missed in a quarter, then the next audit will occur as scheduled for the following quarter. Additional audits may be performed. |
| QAS PM1.0, PM2.5, PM10, TSP, PM2.5 speciation flow audit | Data are suspect back to last AMB staff check. AMS staff must follow up with a flow verification. |
| ## QAS PM2.5, PM10, TSP, PM2.5 speciation AT audit | ## Data are suspect back to last AMB staff check. AMS staff must follow up with an AT verification. |
| ## QAS PM2.5, PM10, PM2.5 speciation FT audit | ## Data are suspect back to last AMB staff check. AMS staff must follow up with an FT verification. |
| ## QAS PM2.5, PM10, TSP, PM2.5 speciation BP audit | ## Data are suspect back to last AMB staff check. AMS staff must follow up with a BP verification. |
| ## QAS PM2.5, PM10, PM2.5 speciation external leak check | ## Data are suspect back to last AMB staff check. PM2.5 and PM10 require an internal leak check. AMS staff must follow up with an external leak check. |
| ## QAS PM2.5, PM10 internal leak check | ## Data are suspect back to last AMB staff check. AMS staff must follow up with an internal leak check. |
| ## QAS PM2.5 speciation leakage test | ## Data are suspect back to last AMB staff check. AMS staff must follow up with a leakage test. |
| ## Filter temperature ˃5 °C above ambient during sampling | ## A QA qualifier “1” is applied to the data. Additional review of data may require a null data qualifier. |

| ## Check | ## Outcome |
| --- | --- |
| ## Sampler run time check | ## Data are invalid for any sample outside 1380 to 1500 minutes. However, when a sample period is less than 1380 minutes, the measured concentration (as determined by the collected PM2.5 mass divided by the actual sampled air volume), multiplied by the actual number of minutes in the sample period and divided by 1440, may be used as if it were a valid concentration measurement for the specific purpose of determining a violation of the NAAQS. Data would then be valid but have a QA qualifier. |
| ## Sampler time check | ## If the time is ˃5 minutes but within 60 minutes of the local standard time, a QA qualifier is applied to the data back to the last known accurate time check. If ˃60 minutes, then a null data qualifier is used back to the last known accurate time check. |
| ## Sampler date check | ## If the date is incorrect on the sampler, data are invalid back to the last known correct date. |
| ## PEP audit comparison results | ## Data are suspect back to last AMB staff check. AMS staff must follow up with a verification. QAS staff may assist. |
| ## Sampler collocation comparison results | ## Data are suspect back to last known accurate concentration. AMS staff and ATS staff must follow up with a thorough review of all processes. QAS staff may assist. |

Table 22. Laboratory Measured Quality Objectives and Outcomes

| Check | Outcome |
| --- | --- |
| Microbalance calibration frequency | If more than 13 months have passed, data will be assigned a QA qualifier “1” if other checks indicate accurate data. |
| Mass reference standards calibration frequency | If more than 13 months have passed, data will be assigned a QA qualifier “1” if other checks indicate accurate data. |
| ## Working vs. primary mass reference standards fail check by ATS staff and QAS staff | ## Further evaluation will determine which weights are inaccurate. Those weights will be sent to the manufacturer for a calibration. |
| ## Filters do not meet integrity check or are mishandled | ## Filters are thrown out and not used. |
| Pre-weighed filters ˃30 days | Filters are thrown out and not used. |

| Check | Outcome |
| --- | --- |
| Inaccurate or missing information on filter sheets | Further review will determine if a QA qualifier or a null data qualifier is needed. |
| PM2.5 and PM10 filter recovery delayed from sampler | A QA qualifier is applied to the data. Additional review of data may warrant a null data qualifier. |
| Clean room exceeds T limits | Filters are invalid if weighed under these conditions and data are assigned appropriate null code qualifier. |
| Clean room exceeds RH limits | Filters are assigned a QA qualifier of “1” if the mean RH is ˃40.4% but ˂41.5% over last 24 hours, humidity control ±5.0% standard deviation over 24 hours, and pre- and post- weighing conditions within ±5% RH; if RH is ≥41.5% over last 24 hours filters are invalid if weighed under these conditions and data are assigned appropriate null code qualifier. |
| Filters not received at clean room within 1 hour after being opened to be logged | A QA qualifier “2” may be applied to the data pending data analysis and review of situation. |
| Exposed filters weighed ˃10 days after sampling | A QA qualifier “1” is applied to the data. Additional review of data may require a null data qualifier |
| Internal QC not performed | Further review will determine if a QA qualifier or a null data qualifier is needed. |
| T above 25.0 °C on post sample transport for PM2.5 and PM10 | Data are invalid and assigned null data qualifier EC. |
| Lab blank ˃ 15 μg | The clean room will be inspected to determine if any contamination is occurring. Review of data will also occur to determine if any impact to samples. |
| PM2.5 and PM10 field blanks ran as a sample | If the field blank is used for a sample makeup, then it will be designated as such. If the field blank ran as a sample due to sampler or operator error, the field blank is invalid and assigned the appropriate null data qualifier. |
| PM2.5 and PM10 field blanks weight gain ˃ 30 μg | Results reported to AMS and QAS. A QA qualifier “1” is applied to the field blank. AMS will investigate sampler cleanliness and additional field blanks may be performed until weight gain is within limits. |
| ## Reweigh of filters ˃ 15 μg | ## Additional reweighs will occur. Further investigation may require the full batch of filters for the day be reweighed. |
| ## Reporting and collocated samples ˃10.0% (PM2.5 and PM10 concentrations ˃ 3 μg/m3 and Pb ˃ .02 μg/m3) | ## ATS will do a reweigh of the PM2.5 and PM10 filters and reanalyze the Pb strips. Further investigation may be needed on field activities to determine any potential issues. |

| ## Check | ## Outcome |
| --- | --- |
| ## Metals solution used for ATS analysis expired | ## Data will have a QA qualifier back to last use of valid solution. Further review of data may require a null data qualifier. |
| ## Speciation filters not conditioned to 4.0 °C or colder after received from sampler and stored at lab | ## A QA qualifier is applied to the data. Additional review of data may require a null data qualifier. |
| ## Speciation filters above 4.0 °C when received at contract laboratory | ## A QA qualifier is applied to the data. Additional review of data may require a null data qualifier. |
| ## QAS reweighs not matching ATS chemist values | ## Additional reweighs are performed until situation is resolved. |
| ## Analysis of QA metals strips by ATS staff is ˃10.0% | ## ATS staff investigates to determine the cause of the difference. Additional strips may be provided to help resolve the situation and as a post check. Field samples back to the last passing QA metals strip analysis may be re-analyzed, have a QA qualifier, or have a null data qualifier. |
| ## QAS clean room audit issues | ## ATS staff will resolve issues and if needed work with the AMS staff to resolve these. |
| ## T and RH sensors in clean room fail QAS staff check | ## Data will be reviewed back to last passing QAS staff check to determine if a QA qualifier or a null data qualifier is needed. |
| ## PM filter cassette fails QAS staff leak check | ## QAS staff will have ATS staff analyze cassette for issues. ATS staff will provide QAS staff with a cassette until the leak check passes. |
| ## Min/Max thermometers for transport/storage of cooler fail QAS staff check | ## The min/max thermometer is not used. |

## Section 15:	Instrument and Equipment Testing, Inspection, and Maintenance Requirements

Samplers and any other items which are part of the data collection process must have checks and routine preventive maintenance performed to ensure proper operation. Most manufacturers supply a preventive maintenance checklist with the instruction manual. When new samplers arrive, a check prior to field use ensures all diagnostics meet manufacturer specifications. The manufacturer normally provides a form showing the QC checks were performed prior to shipping the equipment. All checks and maintenance must be documented in the site logbook, DMDS, or any forms used as part of the check. The AMS parameter specialist is informed of any issues. Any impact to data will be determined on a case-by-case basis unless outlined in this QAPP or in the appropriate AMB SOP. All procedures for performing the diagnostic checks as well as identifying any equipment deficiencies, are outlined in the instruments’ manuals. In cases where equipment does not meet specifications, the AMB staff maintains ready to use spare units to either repair or replace the non-performing equipment. Critical spare parts are also maintained ensuring data collection can continue. Critical spare parts may include items such as spare V-seals. The ATS staff will maintain the clean rooms such that no dust, dirt, trash, etc. could pose a threat for contamination. The ATS staff keep the clean rooms free of dust by sweeping and wiping counters off as needed. ATS staff will also perform a “top-to-bottom” cleaning in March and September.

15.1	PM1.0, PM2.5, PM10 Intermittent and Continuous

Routine preventive maintenance performed by the AMS staff helps prevent failures of the monitoring processes. The overall objective is to increase measurement reliability and prevent data loss. This is accomplished by following and understanding the manufacturer’s recommended guidelines for routine maintenance procedures.

- Filter cassettes – An ATS staff member inspects filter cassettes for contamination after every use. The ATS staff washes the cassettes, and then allows air drying before placing them into plastic bags. QAS staff performs a monthly leak check on filter cassettes in the QA laboratory. ATS staff provides one filter cassette to the QAS staff, who then performs a leak check comparison against a designated QAS filter cassette. Results help eliminate damaged cassette parts and ensure the ATS staff prepares the filter cassettes correctly, which results in no leaks.

- VSCC – An AMS staff member cleans or changes the very sharp cyclone (VSCC) at least once every 30 days for continuous samplers; every 90 days for samplers which run every 3 days; and every 180 days for samplers running every 6 days. More cleaning may occur if needed. The VSCC is used for PM samplers, such as the FRM and FEM for PM2.5. The AMS staff performs maintenance in a separate controlled environment (laboratory, field office, etc.) whenever possible. Exceeding the required time frame can result in data being flagged or a QA qualifier added, pending review of the data. Refer to the Lab Cleaning of Field Items SOP for detailed instructions on maintenance.

- External leak check – Perform an external leak check according to the AMB SOP 2025 Sequential Air Sampler External Leak Check monthly or as indicated by sampler diagnostics.

- Internal leak check – Perform an internal leak check according to the AMB SOP Thermo Fisher Scientific 2025 A-B Sequential Air Sampler Verification and Maintenance when an external leak check fails or as indicated by sampler diagnostics.

- 1st stage inlet – An AMS staff member disassembles the Partisol-Plus Air Sampler sample inlet monthly and cleans with a soft brush or cloth. Exceeding the required time frame can result in data being flagged or a QA qualifier added, pending review of the data. Refer to the Lab Cleaning of Field Items AMB SOP for detailed instructions on maintenance.

- In-line filter – An AMS staff member replaces the in-line filter annually according to the Partisol-Plus Air Sampler operating manual or AMB SOP “2025 A/B Sequential Air Sampler Calibration and Maintenance”.

- V-seals – An AMS staff member checks V-seals once each year or if a leak check fails and replaces seals when necessary, according to the Partisol-Plus Air Sampler operating manual or AMB SOP “2025 A/B Sequential Air Sampler Calibration and Maintenance”.

- Air screens – An AMS staff member cleans the sampler air screens, located under the sampler rain hoods, monthly according to the Partisol-Plus Air Sampler operating manual or AMB SOP “2025 A/B Sequential Air Sampler Verification and Maintenance”.

- Battery voltage – An AMS staff member checks the voltage of the batteries on the main computer board in the electronics compartment annually according to the Partisol-Plus Air Sampler operating manual or AMB SOP “2025 A/B Sequential Air Sampler Calibration and Maintenance”.

- BAM background zero check – This test is performed a minimum of twice per year per manufacturer procedures and the AMB SOP “Met One Instruments Beta Attenuation Monitor (BAM) 1020 Background Zero Check”.

- Solar radiation shield - An AMS staff member inspects the solar radiation shield after each sample run and cleans as necessary with a damp cloth or a dilute soap and water solution. Accumulation of dirt can reduce the effectiveness of the reflective surface and cause the temperature to rise inside the shield.

Maintenance records are kept for each sampler or instrument. The parameter specialist will maintain a logbook for each sampler to track preventive maintenance. If in the field a log is made on the DMDS. These records contain a history of the sampler, including all replacement parts, suppliers, costs, installation dates, etc., (Tables 23 and 24).

# Table 23. Recommended Maintenance and Frequencies

| # Maintenance | # Frequency |
| --- | --- |
| PM1.0 Change the wick | 12 weeks |
| Clean PM10 or PM2.5 inlet | Monthly |
| Large bypass in-line filter exchange | 6 months |
| Battery test | 6 months – change if necessary |
| Pump test | 6 months |
| Clean air inlet system | 1 year |
| Rebuild pump | 1 year – or as needed |

As Needed Maintenance

- Exchange fuses

- Clock adjustment

- Resetting system

- Upgrade system software

Table 24. Inspection of Field Items

| Item | Inspection Frequency | Inspection Parameter | Action if Item Fails Inspection | Documentation Requirement |
| --- | --- | --- | --- | --- |
| Sample downtube | During monthly verification | Visible particulate | Clean with a clean dry cloth | Document in logbook and DMDS |
| VSCC / SCC | During monthly verification | Damaged | Replace | Document in logbook and DMDS |
| Rain collector | Every site visit | ˃1/3 full | Empty | Document in logbook and DMDS |
| Cassette V-seals | As needed | Clean and smooth | Wipe with a clean lint-free dry cloth, or replace as needed | Document in logbook and DMDS |
| In-line filter | Annually | Loaded with particulate | Replace | Document in logbook, DMDS, and calibration form |
| Battery | Annually | Decrease in voltage | Replace | Document in logbook, DMDS, and on calibration form |

15.2	Metals

Routine preventive maintenance helps prevent failures of the monitoring and analytical processes. The overall objective is to increase measurement reliability and prevent data loss. The guidelines below are intended to be general routine maintenance procedures. More detailed information can be found in the manufacturer’s instruction manual for individual instruments as well as the following SOPs.

- Tisch Environmental Hi-Vol+ Lead Sampler Calibration and Maintenance

- Tisch Environmental Hi-Vol+ Lead Sampler Verification and Maintenance

## Maintenance records are kept for each sampler or instrument. The records contain a history of the sampler, including all replacement parts, suppliers, costs, installation dates, etc.

Filter Cassette Maintenance

    - An AMS staff member cleans the filter cassette before any TSP filters are loaded.

Six Month and One Year Maintenance

1.	An AMS staff member replaces motor brushes every six months by following the manufacturer’s instructions. Replacement frequencies vary depending on total operating hours of each motor. A flow controller calibration must be performed by an AMS staff member after any motor maintenance or motor replacement.

2.	An AMS staff member cleans the flow controller probe with alcohol and properly aligns the probe in the throat of the sampler’s cone when needed.

As Needed Maintenance

1.	An AMS staff member checks the gaskets at each setup and replaces them if worn, cracked, or excessively flat.

2.	An AMS staff member cleans the inside of the shelter when needed.

3.	An AMS staff member checks the motor armatures during brush change and replaces when needed.

Laboratory Maintenance and Method Detection Limits for Metals

The ATS staff will perform maintenance on the analysis equipment as stated in the instrument manuals and according to AMB SOP. Part of the maintenance will consist of determining the method detection limits. Definition and procedure for determining the method detection limits are available in 40 CFR Part 136 Appendix B. The definition and procedure are described on the U.S. EPA website, Definition and Procedure for the Determination of the Method Detection Limit, Revision 2. The ATS SOP states the following:

The method detection limit is defined as the minimum measured concentration of a substance that can be reported with 99% confidence that the measured concentration is distinguishable from the method blank results. Analyze a minimum of seven spiked samples and seven blank samples over the course of three or more different non-consecutive calendar days. Results are entered into the Microsoft Excel spreadsheet. The spreadsheet will calculate the standard deviation and will also calculate the method detection limit using the student’s t-value for seven replicates of analysis. Select the greater of method detection limits (standard) or method detection limit blank (Blank) as the initial method detection limit. Method detection limits are determined annually or when there is a significant change in the instrument or procedure.

15.3	Intermittent PM2.5 Speciation

15.3.1	Met One Instruments SASS and Super SASS

The Met One Instruments SASS or Super SASS requires very little maintenance other than the regular checks and calibration activities. Most maintenance involves inspection of various components for damage or wear. Additional information is available in the instrument’s manual and the AMB SOPs Met One Instruments Speciation Air Sampling System (SASS) Calibration and Maintenance and Met One Instruments Speciation Air Sampling System (SASS) Verification and Maintenance. For example:

- An AMS staff member checks the O-rings on a regular basis for signs of deterioration and replaces when necessary. Failure to identify O-ring deterioration may result in leak check or audit flow rate failures.

- An AMS staff member inspects the solar radiation shield after each sample run and cleans as necessary with a damp cloth or a dilute soap and water solution. Accumulation of dirt can reduce the effectiveness of the reflective surface and cause the temperature to rise inside the shield.

- Denuders are replaced by the U.S. EPA contract laboratory staff approximately every three months.

- An AMS staff member cleans the SCC’s grit cap after each sample run and the entire SCC is cleaned or replaced with a clean SCC every three months if running every 3rd day, or every six months if running every six days.

15.3.2	URG 3000N

As with the Met One Instruments SASS or Super SASS, the URG 3000N requires little in the manner of routine maintenance. Additional information is available in the instrument’s manual and the AMB SOPs “URG 3000N Calibration and Maintenance and URG 3000N Verification and Maintenance”.

- An AMS staff member cleans the cyclone every three months if running every 3rd day, or every six months if running every six days.

- An AMS staff member checks the O-rings on a regular basis for signs of deterioration and lubricates with a light coat of vacuum grease if required or replaces as necessary.

- An AMS staff member cleans the interior of the sample and control modules monthly to remove bugs, dirt, and water deposits.

- An AMS staff member cleans the sampler inlet surfaces.

- An AMS staff member checks all Tygon tubing and vacuum lines and replaces as necessary.

- The AMS inspects electrical connections.

- Annually, an AMS staff member cleans the sample inlet tube by pushing a slightly moistened paper towel through the down tube using a wooden dowel rod.

- An AMS staff member and QAS staff inspect filters used for monthly verifications and quarterly audits and replace as needed.

15.4	Continuous PM2.5 Speciation

Aethalometer

Routine preventive maintenance helps prevent failures of the monitoring processes. The overall objective is to increase measurement reliability and prevent data loss. Follow the manufacturer’s recommended guidelines for routine maintenance procedures. Additional information is also available in the AMB SOP “TAPI 633 Aethalometer Verification and Maintenance”.

Maintenance records must be kept for each sampler or instrument. The records should contain a history of the sampler, including all replacement parts, suppliers, costs, installation dates, etc. (see table 25).

# Table 25. Aethalometer Recommended Maintenance

| Maintenance | Frequency |
| --- | --- |
| Optical sampling and analysis cylinder cleaning | 1-2 years as needed |
| Bypass filter cartridge replacement | 1-2 years as needed |
| Sample line | Change out line within 4 years of previous line change unless issues occur |

## As Needed Maintenance

- Exchange fuses

- Adjust clock

- Reset system

- Upgrade system software

## Section 16:	Instrument Calibration and Frequency

All samplers in the air monitoring network adhere to the prescribed calibration schedules which are defined in the sampling methods of 40 CFR Parts 50, 53, and 58; QA GD 2.12; and Quality Assurance Handbook for Measurement Systems Vol. II. Transfer standards used for this work follow strict verification, audit, certification, and calibration schedules, as required by the AMB SOPs listed below, AMB SOPs listed elsewhere in the QAPP, and the AMB QAPP “Calibration, Certification, and Verification Methods of Transfer Standards”.

- Certification and Verification of Flow Devices Using molbox™/molbloc-s™ System

- Certification and Verification of Flow Transfer Standard Using molbox™/molbloc-s™ System

Almost all transfer standards are verified, certified, and/or calibrated against standards which are maintained at the QA laboratory. If QAS staff cannot perform the verification, certification, or calibration, the item is sent to Indiana Standards Laboratory (Indianapolis, Indiana). On occasion some items are sent to Indiana Standards Laboratory to cross check the transfer standards used in the QA laboratory. All equipment is verified, certified, and/or calibrated prior to use. Any equipment used in the particulate program to perform calibrations, verifications, and audits must have documentation kept. A certification file is kept for each item and stored in the QA laboratory and on the AMB shared drive. Although audits are performed with different transfer standards than what is used for a calibration or verification, the audit transfer standards may be checked against the same primary standard used to check calibration or verification transfer standards. Any work performed with an expired transfer standard date is voided. Table 26 lists the transfer standards used in the particulate program.

# Table 26. Calibration and Frequency

| # Item | # Frequency (months) | # Primary Standard | # Limit |
| --- | --- | --- | --- |
| # Temperature probe | # 12 | # Isotech | # ˂ ± 0.51°C |
| # Barometer | # 12 | # Mercury barometer | # ˂ ± 2.26 mmHg |
| # Chinook or flow device | # 12 | # Fluke molbox or Hastings Bubble Meter | # ≤ ± 1% |
| # Orifice | # 12 | # Dresser Roots Meter | # ≤ ± 1% |
| ## Electronic manometer | ## 12 | ## Used with Chinook or flow device on Fluke molbox or Hastings Bubble Meter | # Certified if Chinook or flow device is ## ≤ ± 1% |
| ## S weights | ## 12 | ## Standards laboratory primary weights | ## See EPA QA GD 2.12 |
| ## Pb, Mn, As reference standard solution | ## 12 | ## Manufacturer atomic adsorption certified | ## ≤ ± 1% |

## Section 17:	Inspection/Acceptance Requirements for Supplies and Consumables

Filter media is the main consumable for the particulate network. All PM2.5, PM10, and TSP filters are acquired through the U.S. EPA national contract. PM2.5 and PM10 filter lots are evaluated for stability when received. Stability is performed by ATS randomly picking 9 to 10 filters from different boxes and weighing those filters throughout that year at different times.

All PM2.5 speciation filters are acquired through the U.S. EPA national contract. SASS canisters and URG cartridges with filters and denuders arrive via overland delivery to the AMS in Indianapolis and to the regional offices throughout the state.

Sampler replacement parts, rebuild kits, and continuous particulate tape are obtained from the manufacturer or a distributor when the items are needed by the AMB. The AMS parameter specialist and the ATS staff keep track of their own supplies and order items when needed. The QA laboratory manager ensures the QA laboratory has enough supplies. QA field staff also keep track of when items are needed. A designated QA staff member orders supplies for the QAS when needed. Most consumables do not have an expiration date. If an expiration date exists the AMS parameter specialist, ATS parameters specialist, and the QA laboratory manager track and dispose of expired consumables.

## Section 18:	Non-direct Measurements

IDEM uses all of its own certified data to determine design values and if an area meets the NAAQS (see 40 CFR Part 50). However, non-direct measurement data will be used to support data validation activities. Such data may include historical data or reported concentrations and meteorological measurements from other entities, such as industries and the National Weather Service. Traffic counts from INDOT may also be used (e.g., determine distance air monitoring site must be away from a road). Census population data to determine the appropriate number of samplers/analyzers for an urban area. Data acquired from non-direct measurements may also include site operator observations. The additional data used for site selection and data analysis are viewed as being accurate since this data falls under specific rules and guidelines.

## Section 19:	Data Management

## Particulate data records are generated from a variety of sources in the process of sampling and analyzing of filters. Examples of paper forms are provided in AMB SOPs mentioned throughout this QAPP and are stored on secured AMB computer shared drives or in other secured network drives. The processes of determining valid data and the QC and QA of the processes are mentioned in Section 7 of this QAPP.

## All particulate data is reported in µg/m3 at local conditions except for PM10 (parameter code 81102). The continuous particulate data is reported as hourly averages. The continuous speciation data collected with the Aethalometer provides data as it is collected. However, for some units the actual reported data is collected from the disk monthly and this data replaces the data collected from the instrument.

## The intermittent samples are collected either once or twice per week. The intermittent particulate data consists of 24-hour samples. Specific requirements on collecting and analyzing the data is provided in other sections of this QAPP. All analysis of the particulate samples is performed by ATS staff, except for the intermittent PM2.5 speciation, which is performed by a U.S. EPA contract laboratory staff. The U.S. EPA contract laboratory has its own internal QC and QA checks.

# QA/QC data generated by AMS staff verifications and the QAS staff audits are entered into an audit spreadsheet form. A .pdf copy of the spreadsheet is then uploaded to a shared computer drive. In addition, the data is entered into the U.S. EPA AQS transaction generator resulting in creation of a text file. The flow verification and flow audits are uploaded into AQS quarterly.

## Minimizing data loss is of paramount importance to each monitoring program to meet or exceed the program’s data completeness requirements. Data loss can result from missing or invalid data. Data records are generated through the use of GLIMS for intermittent PM2.5, intermittent PM10, and intermittent PM10c; DMDS for continuous PM1.0, continuous PM2.5, continuous PM10, continuous PM10c, and continuous PM2.5 speciation; and an Excel spreadsheet for intermittent TSP for metals. Data records for intermittent PM2.5 speciation are handled through a U.S. EPA contractor. The goal of IDEM is to collect 100% of generated data and maintain a completeness rate of least 75% valid data. The U.S. EPA requires at least 75% of data to be valid to meet completeness requirements. The process for determining valid data and the QC and QA of these processes is mentioned in this QAPP. The data process includes collecting, storing, transmitting, verifying, validating, and reporting to the U.S. EPA’s AQS database.

## The Ambient particulate data must be collected, verified, validated, and reported to the U.S. EPA’s AQS database quarterly and certified yearly. Quarterly data is reported to the AQS within 90 days after the quarter is completed. Data consists of sample concentrations, flow verifications, and flow audits. Continuous particulate data is made available to the IDEM website as it is collected. Therefore, the data has not initially had QC or QA checks performed. Continuous particulate data is also provided to AIRNOW as it is being collected. The process for determining valid data and the QC and QA of these processes is mentioned in this QAPP.

## The data process starting from the initial check of the filter media to the field, to AQS submittal depends upon the parameter. Each type of particulate parameter is described below.

## 19.1 Intermittent PM2.5, Intermittent PM10, and Intermittent PMc

- Sample and data collection are performed using an approved U.S. EPA reference sampler.

- The sampler is calibrated prior to sample collection. Information is stored on an AMS calibration sheet and logged in DMDS.

- Filters to be sampled go through a candling process. Filters with holes or scratches are discarded. The process is documented in filter conditioning log, blanks log, and room maintenance binder.

- Filters are conditioned for at least 24 hours. The process is described in the SOP “Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software”. The date and time when the filters could be weighed is recorded on a piece of tape on the tray on which the filters were placed.

- Filters are labeled. The process is described in the AMB SOP mentioned in step 4.

- The tare weight of the filter is measured and recorded in GLIMS.

- The QAS staff also performs a tare re-weigh on at least 7% of the batch of filters, but no less than three filters. The QA audit is documented in an ATS Excel spreadsheet, which is password protected.

- Filters are packed in magazines and placed in the appropriate coolers. The process is documented in the AMB SOP mentioned in step 4.

- Samples are set up prior to the sample day. Information is documented on a data sheet, a site logbook, and in DMDS.

- Samples are picked up after sampling is completed. Information is documented on a data sheet, a site logbook, and in DMDS.

- Samples are brought back to the office and logged in GLIMS. The min/max temperatures are recorded. The information is recorded on the data sheet.

- Samples are taken to the clean room and unpacked from the magazine by an ATS staff member. Information is documented on the data sheet.

- Samples are conditioned for at least 24 hours. The process is described in the AMB SOP mentioned in step 4.

- The gross weight is measured. Information is documented on the weighing conditions spreadsheet; the HOBO log, which is the ATS temperature and relative humidity clean room data logger; and the clean room logbook.

- Once all laboratory functions are complete for a filter, the samples are archived in cold storage for 12 months as stated in the AMB SOP “Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software”. After 12 months filters are stored in the AMB warehouse for a minimum of 5 years.

- Data are screened by AMS staff monthly per AMB SOP “Filter Based Particulate Matter Data Validation”. Data sheets are filed by day 15 of the following month and verification completed by day 30.

- Data are sent to the QAS program coordinator to validate within 15 working days per AMB SOP “PM2.5/10 Intermittent Data Check Procedures”. Information is documented on a data check sheet and stored on the AMB computer shared drive. Any exceedance review is performed within 15 days after the validation process is completed.

- The AMS AQS administrator is informed by the QAS program coordinator that data validation is complete.

- The AMS AQS administrator uploads validated data into AQS.

- Monthly AMS staff verifications and quarterly QAS staff audits are performed. The results are entered into the U.S. EPA transaction generator and the resulting transaction file is then created so it can be upload to AQS.

- Once all data for a quarter are verified and validated, a designated QAS environmental manager submits QA flow audits into AQS, and an AMS environmental manager submits monthly flow verifications into AQS.

- Each quarter AMP reports are submitted by a QAS environmental manager to the QAS chief and are reviewed for accuracy and any issues. Any indication of a problem requires further analysis, by QAS staff, ATS staff, or AMS staff, of the reported data with the possibility of data resubmission.

- Each year PM2.5, PMc and PM10 data from the previous year is certified as being true and accurate. The AMB chief, AMS (1 and 2) chiefs, ATS chief, QAS chief, and specific AMS (1 and 2) staff and QAS perform a final check on this data package, which is then submitted to U.S. EPA.

## 19.2 Intermittent PM2.5 Speciation

- Sample and data collection are performed using an approved U.S. EPA reference sampler.

- The sampler is calibrated prior to sample collection. Information is stored on an AMS calibration sheet and logged in DMDS.

- Samples are set up prior to the sample day. Information is documented on a data sheet, a site logbook, and in DMDS.

- Samples are picked up after sampling is completed. Information is documented on a data sheet, a site logbook, and in DMDS.

- Samples are brought back to the office and stored in a refrigerator prior to shipping based on the U.S. EPA shipping schedule. Temperature of the storage area is documented to ensure filters meet requirements.

- Filters are shipped to the U.S. EPA contract laboratory as scheduled. Once analysis is performed, the AMS staff receives monthly batch data from the U.S. EPA contractor via Data Analysis Reporting Tool (DART). The data are analyzed by the AMS staff per the AMB SOP “Intermittent Chemical Speciation Data Validation”. Changes to data can occur based on the AMS staff findings. The U.S. EPA contractor submits the data into AQS.

- Monthly AMS staff verifications and quarterly QAS staff audits are performed. The results are entered into the U.S. EPA transaction generator and the resulting transaction file is then created so it can be upload to AQS.

- Once all data for a quarter have been verified and validated, a designated QAS environmental manager submits QA flow audits into AQS, and an AMS environmental manager submits monthly flow verifications into AQS.

- Each quarter AMP reports are submitted by a QAS environmental manager to the QAS chief. The reports are reviewed for accuracy and any issues. Any indication of a problem requires further analysis, by QAS or AMS staff, of the reported data with the possibility of data resubmission.

## 19.3 Intermittent TSP for Metals

- Sample and data collection are performed using an approved U.S. EPA reference sampler.

- The unit is calibrated prior to sample collection. This information is stored on an AMS calibration sheet and logged in DMDS.

- Samples are set up prior to the sample day. Information is documented on a filter card and in DMDS.

- Samples are picked up after sampling is completed. Information is documented on a filter card and in DMDS.

- The filters are given to the AMS parameter specialist, who records data from the filter card into an Excel spreadsheet, which is stored on the AMB shared drive. The AMS environmental manager also inspects the filter for damage and reviews all the data to determine whether the filter is valid.

- The AMS environmental manager gives the filter to the QAS program coordinator, who does a QA check on the sample per the AMB SOP “Particulate Filter Quality Assurance”. The filter is then given to the ATS for analysis.

- The filters are extracted and analyzed by ATS environmental chemist per AMB SOP “Atomic Absorption Spectrometer Analysis for Lead and other Trace Metals”.

- Once analysis is completed by the ATS staff, the data are added to an Excel spreadsheet, which is located on the AMB shared drive.

- The AMS staff and the QAS program coordinator reviews the data as per the AMB SOP “Metals Total Suspended Particulate (TSP) Intermittent Data Review”.

- Monthly AMS staff verifications and quarterly QAS staff audits are performed. The results are entered into the U.S. EPA transaction generator and the resulting transaction file is then created so it can be upload to AQS.

- Once all data for a quarter are verified and validated, a designated QAS environmental manager submits QA flow audits into AQS, and an AMS environmental manager submits monthly flow verifications into AQS.

- Each quarter AMP reports are submitted by a QAS environmental manager to the QAS chief and are reviewed for accuracy and any issues. Any indication of a problem requires further analysis, by QAS or AMS staff, of the reported data with the possibility of data resubmission.

- Each year TSP-Pb data from the previous year is certified as being true and accurate. The AMB chief, AMS (1 and 2) chiefs, ATS chief, QAS chief, and specific staff in the AMS (1 and 2) and QAS perform final check on this data package, which is then submitted to U.S. EPA.

## 19.4 Continuous PM1, Continuous PM2.5, Continuous PM10, and Continuous PMc

- Sample and data collection are performed using an approved U.S. EPA reference analyzer.

- The unit is calibrated prior to sample collection. Information is stored on an AMS calibration sheet and logged in DMDS.

- Data is sent from the analyzer digitally to the on-site data logger. Although data collection is continuous, data is averaged in five-minute intervals. The data logger can store up to approximately two weeks of data before it starts being overwritten. The data logger at the site requires a password to access it.

- The DMDS communicates with the data loggers and retrieves data every 15 minutes through a cellular modem. The DMDS software collects data from the data logger using the Native Datalogger Computer-To-Computer language (CC-SAIL). Preliminary data checks are performed automatically before data is stored on the DMDS server.

- The data is decoded, checked for errors, processed, and stored in a database. The DMDS serve has a mirror backed for disaster recovery. The raw data can also be reloaded if a needed, such as accidental deletion of processed numbers or erroneous slope, intercept, offset applied to data.

- The DMDS software performs quality control checks on the span/zero checks (SPZ), span/1-point QC/zero checks (SPN) and calibrations (CAL) to determine if they meet tolerance limits. The results of the automatic quality checks can result in data flagging, affecting data validity.

- Each business day, the AMS DMDS Administrator will check the monitoring data database to ensure that all data were polled and transmitted successfully from each monitoring station and stored on DMDS. In case of missing data, a re-poll can be initiated to backfill the data.

- Approximately one to four weeks after the end of a month, AMS staff will initiate the data verification procedure. AMS staff review and flag raw data and document data verification within the DMDS. The DMDS requires a log entry when any changes are made to the data. The logs document who verified the data, changes made, and a date/time when this occurred (see AMB SOP “Gaseous Data Validation Using DMDS”).

- Once the data has been verified, AMS staff will inform the QAS program coordinator via e-mail that verified data is ready for validation.

- The QAS program coordinator logs the date that the verified data has been sent to the QAS and then informs QAS environmental manager via e-mail that data is ready for validation. The QAS environmental manager has 15 business days to complete the validation process. Any exceedance review is performed within 15 days after the validation process is completed.

- The QAS environmental manager logs into the DMDS server (read access only) and performs an audit on the data (validation process) using the various DMDS reports. The validation process is documented on a Validation Check form, which is stored on the Branch shared drive.

- If data issues arise during the validation process, the issues are sent to the appropriate AMS staff member for correction or additional information.

- Once the validation review is complete, the QAS environmental manager will electronically initial and date the Validation Check form. The QAS environmental manager informs the QAS program coordinator via e-mail that the data validation is complete.

- The QAS program coordinator will log the data validation completion date and informs the AMS AQS administrator via email that the data review process is completed by the QAS.

- The AMS AQS administrator downloads the validated data from the DMDS and submits the data to AQS.

- Monthly AMS staff verifications and quarterly QAS staff audits are performed. The results are entered into the U.S. EPA transaction generator and the resulting transaction file is then created so it can be upload to AQS.

- Once all data for a quarter are verified and validated, a designated QAS environmental manager submits QA flow audits into AQS, and an AMS environmental manager submits flow-rate verifications into AQS.

- Each quarter AMP reports (251, 256, 450) are submitted by a QAS environmental manager to the QAS chief, and these are reviewed for accuracy and any issues. Any indication of a problem will require further analysis by QAS and/or AMS of the reported data with the possibility of data resubmission.

- Each year PM2.5, PMc, and PM10 data from the previous year is certified as being true and accurate. The AMB chief, AMS (1 and 2) chiefs, ATS chief, QAS chief, and specific staff in the AMS (1 and 2) and QAS do a final check on this data package, which is then submitted to U.S. EPA.

## 19.5 Continuous PM2.5 Speciation

- Sample and data collection are performed using an Aethalometer.

- The Aethalometer is calibrated prior to sample collection. The information is stored on an AMS calibration sheet and logged in DMDS.

- Data is sent from the analyzer digitally to the on-site data logger. Although data collection is continuous, data is averaged in five-minute intervals. The data logger can store up to approximately two weeks of data before it starts being overwritten. The data logger at the site requires a password to access it.

- The DMDS communicates with the data loggers and retrieves data every 15 minutes through a cellular modem. The DMDS software collects data from the data logger using the Native Datalogger Computer-To-Computer language (CC-SAIL). Preliminary data checks are performed automatically before data is stored on the DMDS server.

- The data is decoded, checked for errors, processed, and stored in a database. The DMDS serve has a mirror backed for disaster recovery. The raw data can also be reloaded if a needed, such as accidental deletion of processed numbers or erroneous slope, intercept, offset applied to data.

- The DMDS software performs quality control checks on the span/zero checks (SPZ), span/1-point QC/zero checks (SPN) and calibrations (CAL) to determine if they meet tolerance limits. The results of the automatic quality checks can result in data flagging, affecting data validity.

- Each business day, the AMS DMDS Administrator will check the monitoring data database to ensure that all data were polled and transmitted successfully from each monitoring station and stored on DMDS. In case of missing data, a re-poll can be initiated to backfill the data.

- Once the data are verified, AMS staff informs the QAS program coordinator via e-mail that verified data are ready for validation.

- The QAS program coordinator logs the date that the verified data was sent to the QAS staff and then informs the QAS environmental manager via e-mail the data are ready for validation. The QAS environmental manager has 15 business days to complete the validation process.

- The QAS environmental manager takes and loads the data, which is a .txt file, into a macro Excel spreadsheet, which allows for easier viewing. Data are screened for high and low values, audits, and verifications; and confirms existing null data codes. Also, the monthly logs in DMDS are checked to confirm work performed at the site. The validation process is documented on a Validation Check form, which is stored on the branch shared drive.

- If data issues arise during the validation process, the issues are sent to the appropriate AMS staff member for correction or additional information.

- Once the validation review is complete, the QAS environmental manager electronically initials and dates the Validation Check form. The QAS environmental manager sends the data to the AMS AQS administrator and cc’s the QAS program coordinator that the data validation is complete.

- The QAS program coordinator logs the data validation completion date.

- The AMS AQS administrator submits the validated data to AQS.

- Monthly AMS verifications and quarterly QAS audits are performed. The results are entered into the U.S. EPA transaction generator and the resulting transaction file is then created so it can be upload to AQS.

- Once all data for a quarter has been verified and validated, a designated QAS environmental manager submits QA flow audits into AQS, and an AMS environmental manager submits flow-rate verifications into AQS.

- Each quarter AMP reports are submitted by a QAS environmental manager to the QAS chief and are reviewed for accuracy and any issues. Any indication of a problem requires further analysis, by QAS or AMS staff, of the reported data with the possibility of data resubmission.

## Section 20:	Assessments and Response Actions

## IDEM utilizes several assessment procedures to ensure activities are being conducted as planned and data are acceptable. This also helps to identify and correct issues. These procedures may include formal or informal communication and response (Table 27).

Table 27. Assessments and Response Actions

| ## Assessment | ## Conducted By | ## Frequency | ## Requirement |
| --- | --- | --- | --- |
| ## Monthly flow rate verification | ## AMS staff | ## Monthly | ## Tables 10 and 12 of this QAPP |
| Semi-annual flow rate audit | QAS staff | Quarterly for each sampler | ## Tables 11 and 13 of this QAPP |
| Precision via permanently collocated sampler | AMS staff | Every 6 days | ## ˂10.1 % CV for total PM precision; ˂20.1 %CV for total Pb precision |
| Data verification | AMS staff | Monthly | ## Screening of data sample information, QC checks |

| Assessment | Conducted by | Frequency | ## Goals |
| --- | --- | --- | --- |
| Data validation | QAS staff | Monthly | ## Screening of data, sample information, QC and QA checks |
| ## Site Evaluation | ## QAS staff | ## 2 years per site | ## QAPP requirements |
| 5-year network assessment | AMB staff | 5 years | ## 40 CFR § 58.10 |
| PEP | U.S. EPA Regional staff or contractor | See Section 7.2 of this QAPP | ## Section 7.2 of this QAPP |
| Internal clean room audit | QAS staff | Annually | ## U.S. EPA QA GD 2.12 |
| Technical systems audit | U.S. EPA Regional staff | Once every 3 years | ## 40 CFR 58 ## Appendix A § 2.5 |
| Exceedance report | QAS staff | As needed | ## Screening of information, such as calibration, pickup date, analysis date, etc., which may confirm the high value |
| AMP reports | QAS staff | Quarterly | ## Review data for statistical issues and completeness |
| QAPP | QAS staff | Annually | ## Determine if changes are needed which accurately describes the project |
| ANP | AMB staff | Annually | ## 40 CFR § 58.10 |
| Annual data certification | AMB staff | Annually | ## 40 CFR § 58.15 |

## In the event an assessment identifies an area of concern, specific corrective actions occur depending on what the finding reveals. Listed below are the assessments and corrective action time frames for follow-up.

## Monthly verification – AMS staff address issue once made aware. QAS environmental manager ensures action was taken and issues were resolved.

## Semi-annual flow rate audit – QAS staff address issues either from the site or once back in the office. Consult with AMS parameter specialist to resolve issues.

## Precision via permanently collocated sampler – QAS staff addresses issues once CV is determined. AMS 1 chief and ATS chief are informed of potential issues. A QA qualifier or null data qualifier may be applied to the data and procedures may need to be adjusted.

## Data verification process – The AMS parameter specialist performs a verification of the ambient data within 1 to 3 weeks after the end of a month. Documentation of the process can be found in the DMDS validator notes.

Data validation process – A QAS staff member performs a validation review of the data within 15 days after the QA program coordinator is informed that a data package is available for review. Data validation is documented on the Data Check Sheets, which are stored on the AMB shared drive.

Exceedance report – A QAS staff member performs additional reviews on any high values which meet or exceed the NAAQS for the parameter. The review is performed within 15 days after the data validation is completed for the specific site and month. The exceedance is documented on a memo, which is stored on the AMB shared drive.

AMP reports – Within 90 days after the end of the quarter, quarterly AMP Reports (AMP256, AMP430) are sent to the QAS chief after the flow verifications and PE audit results are submitted to AQS. The QAS chief and an environmental manager may consult with AMS staff or ATS staff to resolve any issues once any are discovered. If practical, the issue should be resolved within two weeks after identification.

## QAPP - QAS chief ensures QAPP is followed and makes necessary changes, with approval by AMB chief, AMS (1 and 2) chiefs, and ATS chief. An annual check is documented or when a change needs to be made in the QAPP revision history.

## ANP - The annual network plan is due to U.S. EPA Regional Administrator by July 1. The AMB tries to have a complete ANP available for public comment by mid-May to allow for a 30-day public comment period which is completed by mid-June. Any comments from the public will be addressed before final submittal to U.S. EPA. Once U.S. EPA does their review, if needed, AMB will provide any feedback prior to the new year. Depending on what type of changes may be needed after the U.S. EPA reviews the ANP, corrective actions will be taken immediately, and projected to be completed prior to the new year.

## Annual data certification – The annual data certification for ambient data from the previous year (January 1 – December 31) is due to the U.S. EPA Regional Administrator by May 1st of each year. AMB staff review the annual data package and resolve any correctable issues, if practical, prior to submission of the certification package.

## Site Evaluation – QAS staff perform site evaluations on a 2-year cycle. AMS staff address issues based on QAS findings. Issues are usually resolved within a business day when data are impacted; or some time frames may be extended based on the nature of the issue and whether the site is on private property. QAS chief ensures corrective action is taken to resolve the issues.

## 5-year network assessment – Due at U.S. EPA by July 1 for years ending in zero or five. Corrective actions by AMB staff based on U.S. EPA feedback within the time frame allotted for the response.

## PEP – QAS chief works with AMS staff to address PEP issues as soon as possible, usually within a week to resolve any issues.

## Internal clean room audit – QAS chief works with ATS staff  to address issues as soon as possible.

## Technical Systems Audit - The U.S. EPA Regions staff schedule on a 3-year frequency. The QAS staff works with AMS staff and ATS staff to address any audit findings. All findings should be resolved within 1 year of the TSA report.

Section 21:	Reports to Management

## Reports that are generated and utilized in the particulate program are listed in table 28.

Table 28. Reports to Management

| Report | Frequency | Responsible Party |
| --- | --- | --- |
| PM2.5 exceedance report | Monthly | AMS |
| Pb design value | Monthly | AMS |
| AMP reports | Quarterly | QAS |
| Primary vs. working standard weight comparison | Quarterly | ATS |
| Chemist weighing comparisons | Annually | ATS |
| PM2.5, PM10, and Pb exceedance reports | As needed | QAS |
| PM2.5 and Pb CV reports | Quarterly | QAS |
| Invalid data memos | As needed | AMB |
| Annual network plan | Annually | AMB |
| Annual Data Certification Package | Annually | AMB |
| 5-year network plan | 5 Years | AMB |

## Section 22:	Data Validation and Usability

Many criteria used to review and validate data are detailed in the previous sections of this QAPP. The AMB staff utilizes this established QAPP, SOPs, and U.S. EPA Validation Template found in Appendix D of “Quality Assurance Handbook for Air Pollution Measurement Systems, Volume II: Ambient Air Quality Monitoring Program", to determine data validity. Once data has gone through all the processes, then it is deemed usable for its intended objective.

## Section 23:	Validation and Verification Methods

Data verification is the process of evaluating the completeness, correctness, and conformance of a specific data set against the method, procedural, or contractual requirements, as specified in both the SOPs and 40 CFR Part 58. Data validation is a process which extends the evaluation of data beyond method, procedural, or contractual compliance (i.e., data verification) to ensure that reported values meet the quality goals of the environmental data operations and that the data can be used for the intended purpose.

The AMB staff uses the criteria particulate pollutant validation templates provided in Appendix D of the U.S. EPA QA Handbook for Air Pollution Measurement Systems: Volume II: Ambient Air Quality Monitoring Program (EPA-454/B-17-001, January 2017) for the weight of evidence approach for validating criteria particulate data. The AMB staff follows the guidance in the QA Handbook regarding the use of the templates and handles the criteria as follows:

- Critical criteria are issues deemed critical to maintaining the integrity of the hourly ambient concentration measurement or a group of successive hourly ambient concentration measurements. Data reviewers should invalidate observations which do not meet all criterion in the critical criteria table unless compelling reasons and justification exist for not doing so. The hourly measurement or group of hourly measurements which do not meet one or more of these criteria is invalid unless proven otherwise. In most cases, the CFR dictates the requirement, the implementation frequency of the criteria and the acceptance criteria so these criteria are considered regulatory in nature.

- Operational criteria are situations where violations of a criterion or criteria may be cause for invalidation of the data. Data reviewers should consider other QC information which may or may not indicate the data are acceptable for the parameter they want to control. Therefore, ambient data, which do not meet one or more of these criteria, are suspect unless other QC information demonstrates otherwise, and the reviewers have adequate documentation of the information. Data reviewers should investigate, mitigate, or justify the reason for not meeting the criteria.

- Systematic criteria include criteria, including the DQOs, which are important for the correct interpretation of the data, but do not usually impact the validity of the ambient data. If the data do not meet the DQOs, this does not invalidate any of the hourly measurements, but it may impact the confidence in the attainment or non-attainment decision.

The AMB staff brackets all particulate concentration data using the results of the monthly verifications, calibration, or performance evaluation audit to ensure the particulate samplers were in proper operating condition between the checks. When an intermittent sampler fails a check, data are suspect back to the last passing check until the cause of the failure is remedied. For some continuous samplers, the DMDS software automatically flags the data when the error occurs to time and date when the issue causing failure is remedied. The AMS parameter specialist reviews the data, determines the cause for the failure, and verifies the extent of the data invalidation period. During the validation process, the QAS staff reviews the invalid data period to ensure data are proper, accurate, and documented.

23.1 AMS Verification for Intermittent PM2.5, Intermittent PM10, and Intermittent PMc

The verification on intermittent filters begins when the ATS staff performs the initial weighing and again after the filters have sampled as per the AMB SOP “Analysis of Both PM2.5 and PM10 Particulate Matter Using GLIMS Software”. Once the concentrations are determined, the AMS staff performs the data verification as per the AMB SOP “Filter Based Particulate Matter Data Validation”. Data flags can be added during the sampling process, the weighing process, and after the concentrations are determined and entered into GLIMS. Any flags to be added can be found in Table 29 of this QAPP and on the U.S. EPA website,  Qualifiers - AQS Reference Table.

23.2 AMS Verification for Intermittent PM2.5 Speciation

The AMS parameter specialist receives monthly batch data from the U.S. EPA contract laboratory, via DART. The data have had QC checks performed by the U.S. EPA contract laboratory. The AMS environmental manager then performs data analysis per the AMS SOP “Intermittent Chemical Speciation Data Validation”. Any flags to be added can be found in Table 29 of this QAPP and on the U.S. EPA website,  Qualifiers - AQS Reference Table.

23.3 AMS Verification for Intermittent TSP for Metals

The verification on TSP for metals filters begins when the ATS staff performs the initial candling and again after the filters have sampled. The AMS staff also reviews data collected for the sample and scans the filter for defects once the filter has been retrieved from the sampler and brought back to the office. Once the concentrations are determined, the ATS staff and the AMS staff determine whether the concentrations are accurate. Data flags can be added during the sampling process, the analysis process, and after the concentrations are determined and entered in an Excel spreadsheet. Any flags to be added can be found in Table 29 of this QAPP and on the U.S. EPA website, Qualifiers - AQS Reference Table.

23.4 AMS Verification for Continuous PM1.0, Continuous PM2.5, Continuous PM10, and Continuous PMc

The DMDS administrator reviews the DMDS reports every business day to check for anomalies and to repoll data loggers when missing data are notated. The DMDS administrator reviews the hourly values for any NAAQS exceedances and notifies the AMS parameter specialist and the QAS staff. After a month of data are collected, a monthly pollutant concentration report is generated by DMDS. The AMS staff verifies all data per the AMB SOP “Continuous Particulate Data Validation” during the monthly data review. The monthly report is reviewed by the AMS parameter specialist for data values as well as any flags applied by DMDS. The review uses the DMDS parameter specialist to justify the application of the flag and to determine whether flags are accurately applied. The data traces are also evaluated. Any changes needed are applied by the AMS parameter specialist. Once the process is completed the AMS parameter specialist leaves an “Annotation Log” in DMDS, which means the data are verified. Once the verification process is completed by the AMS staff, the QAS program coordinator is informed by the AMS parameter specialist that the data are ready for validation.

The null and data qualifiers listed in Table 29 are available to flag data as appropriate. The DMDS automatically applies a flag text to most. The flag text is then changed to the appropriate qualifier.

Table 29. Null and Data Qualifiers

| Qualifier Description | NULL Data | Data Qualifier | Qualifier Type |
| --- | --- | --- | --- |
| Sample Pressure out of Limits. | AA |  |  |
| Technician Unavailable. | AB |  |  |
| Construction/Repairs in Area. | AC |  |  |
| Shelter Storm Damage. | AD |  |  |
| Shelter Temperature Outside Limits. | AE |  |  |
| Scheduled but not Collected. | AF |  |  |
| Sample Time out of Limits. | AG |  |  |
| Sample Flow Rate or CV out of Limits. | AH |  |  |
| Insufficient Data (cannot calculate). | AI |  |  |
| Filter Damage. | AJ |  |  |
| Filter Leak. | AK |  |  |
| Voided by Operator. | AL |  |  |
| Miscellaneous Void. | AM |  |  |
| Machine Malfunction. | AN |  |  |
| Bad Weather. | AO |  |  |
| Vandalism. | AP |  |  |
| Collection Error. | AQ |  |  |
| Lab Error. | AR |  |  |
| Poor Quality Assurance Results. | AS |  |  |
| Calibration. | AT |  |  |
| Monitoring Waived. | AU |  |  |
| Power Failure. | AV |  |  |
| Wildlife Damage. | AW |  |  |
| Q C Audit. | AZ |  |  |
| Maintenance/Routine Repairs. | BA |  |  |
| Unable to Reach Site. | BB |  |  |
| Multi-point Calibration. | BC |  |  |
| Building/Site Repair. | BE |  |  |
| Lost or damaged in transit. | BI |  |  |
| Operator Error. | BJ |  |  |
| Site computer/data logger down. | BK |  |  |
| QA Audit. | BL |  |  |
| Exceeds Critical Criteria. | EC |  |  |
| Qualifier Description | NULL Data | Data Qualifier | Qualifier Type |
| Filter Inspection Flag. | FI |  |  |
| Storm Approaching. | SA |  |  |
| Sampler Contamination. | SC |  |  |
| Sample Volume out of limits. | SV |  |  |
| Holding Time Or Transport Temperature Is Out Of Specs. | TS |  |  |
| Deviation from a CFR/Critical Criteria Requirement. |  | 1 | QA |
| Data reviewed and validated. |  | 1V | QA |
| Operational Deviation. |  | 2 | QA |
| Field Issue. |  | 3 | QA |
| Lab Issue. |  | 4 | QA |
| Outlier. |  | 5 | QA |
| QAPP Issue. |  | 6 | QA |
| Field Blank Value Above Acceptable Limit. |  | FB | QA |
| Filter Integrity Issue. |  | FX | QA |
| Sample pick-up hold time exceeded. |  | HT | QA |
| Influenced by nearby source. |  | NS | QA |
| Pressure Sensor Questionable. |  | QP | QA |
| Temperature Sensor Questionable. |  | QT | QA |
| Does not meet QC criteria. |  | QX | QA |
| Value substituted from secondary monitor. |  | SS | QA |
| Does Not Meet Siting Criteria. |  | SX | QA |
| Trip Blank Value Above Acceptable Limit. |  | TB | QA |
| Transport Temperature is Out of Specs. |  | TT | QA |
| Validated Value. |  | V | QA |
| Value below normal; no reason to invalidate. |  | VB | QA |
| Flow Rate Average out of Spec. |  | W | QA |
| Filter Temperature Difference or Average out of Spec. |  | X | QA |
| Elapsed Sample Time out of Spec. |  | Y | QA |
| African Dust. |  | IA | INFORM |
| Asian Dust. |  | IB | INFORM |
| Chem. Spills & Industry Accidents. |  | IC | INFORM |
| Cleanup After a Major Disaster. |  | ID | INFORM |
| Demolition. |  | IE | INFORM |
| Fire - Canadian. |  | IF | INFORM |
| Fire - Mexico/Central America. |  | IG | INFORM |
| Fireworks. |  | IH | INFORM |
| Qualifier Description | NULL Data | Data Qualifier | Qualifier Type |
| High Pollen Count. |  | II | INFORM |
| High Winds. |  | IJ | INFORM |
| Infrequent Large Gatherings. |  | IK | INFORM |
| Other. |  | IL | INFORM |
| Prescribed Fire. |  | IM | INFORM |
| Seismic Activity. |  | IN | INFORM |
| Stratospheric Ozone Intrusion. |  | IO | INFORM |
| Structural Fire. |  | IP | INFORM |
| Terrorist Act. |  | IQ | INFORM |
| Unique Traffic Disruption. |  | IR | INFORM |
| Volcanic Eruptions. |  | IS | INFORM |
| Wildfire-U. S. |  | IT | INFORM |
| Construction. |  | J | INFORM |

Additional flags are available for application as well. The flags are found at https://aqs.epa.gov/aqsweb/documents/codetables/qualifiers.html.

23.5 AMS Verification for Continuous PM2.5 Speciation

The DMDS administrator reviews the DMDS reports every business day to check for anomalies and to repoll data loggers when missing data are notated. The AMS staff verifies all data per the AMB SOP “Aethalometer Data Validation Procedures” during the monthly data review. The monthly report is reviewed by the AMS parameter specialist for data values as well as any flags applied by the masher program. Once the verification process is completed by the AMS staff, the QAS program coordinator is informed by the AMS parameter specialist that the data are ready for QA.

23.6 QAS Validation

The QAS staff validates the continuous particulate data per the AMB SOP “DMDS Validated Data Review Procedures” and the intermittent particulate data per QAS SOP “Particulate Filter Quality Assurance”. The QAS program coordinator sends an e-mail to the QAS staff responsible for performing the validation on the specific parameter, who then has 15 working days to complete the data check. A validation form is used by QAS staff, which includes analysis of concentrations and a review of QC and QA processes, and to document the review.

When high concentrations are identified whether by email or memo by the AMS staff or ATS staff, the QAS staff is notified of the site, and date and time when it occurred. In some instances, a high concentration could be discovered during the data audit process. In either situation, the QAS staff initiates a NAAQS, exceedance memoranda. A standard form is used to identify whether the data exceeding the NAAQS is of a valid nature. The exceedance verification form identifies results of specific checks preceding and subsequent to the exceedance event. If a check on both sides of the NAAQS exceedance are considered valid, then the data are considered valid provided that no intervening null codes between the checks invalidates the data. If any checks indicate an issue, the data may be suspect. Checks may be quarterly QAS PE audits, AMS staff monthly verification checks, or ATS staff reweighs. In some situations, this may entail a special QA audit after the exceedance to verify analyzer performance and data validity. Any invalid or missing data may be documented in the comments portion of the exceedance form.

The QAS chief and one QAS environmental manager reviews various AMP reports (e.g., AMP251, AMP256, AMP350, and AMP450) generated quarterly from AQS for any data which requires further analysis. The QAS staff has the authority to have data rechecked and if needed, invalidate additional data.

## Section 24:	Reconciliation with Data Quality Objectives

## The data quality objectives and intended uses for the particulate pollutant data are discussed in Section 7 of this QAPP. The main purpose of the data are to show compliance with the U.S. EPA NAAQS and to measure air pollutant concentrations which may be of concern for public health and public welfare considerations. Section 7 of this QAPP also lists the measurement quality objectives, which were established to provide the expected data quality users require.

## The role of this QAPP is to establish procedures to control measurement uncertainty to an appropriate level to achieve the objectives for which monitoring data are collected. If guidelines, any SOPs, and any QAPPs governing the measurement process are followed and all measurement quality objectives listed in this QAPP are met, the DQOs can be achieved. However, a chance always exists where exceptional field events may negatively affect the performance of the monitoring station. Therefore, reconciling the monitoring data with the DQOs is important to evaluating whether the data set is adequate for its intended use. The evaluation involves reviewing routine data, such as the monthly verification and validation reviews described in Section 23 of this QAPP, and the results of monthly verification checks.

## On a quarterly basis, the performance of the monitoring network will be evaluated by reviewing the data quality statistics (precision, bias, and completeness) of the QA and QC data set and comparing the results to the monitoring project’s goals. Data quality assessment statistics are taken from the AQS AMP450 Report Quick-look Report, the AQS AMP430 Report Data Completeness, and the AQS AMP256 Report Data Quality Indicator Report. The AMP450 Report provides summary statistics on the criteria pollutant data collected. The AMP430 Report provides a status of the quantity of criteria pollutant data collected. The AMP256 Report provides a status of the QA and QC activities. Unacceptable performance for any of the DQO goals does not automatically indicate that the data set cannot be used for its intended purpose, i.e., the support of the NAAQS attainment decision process. However, the impact on the confidence with which the data set can be used for its intended purpose in the decision process will have to be reviewed and communicated. This is done in the quarterly reports generated by the QAS environmental manager and QAS chief. Any anomalies are reported to the AMSs and ATS chiefs, and the AMB chief. The reports identify the points at which the data failed to meet DQOs and at what point in time, after corrective action, the data again meet DQOs. The corresponding data are flagged and commented, and all supporting documentation is included in the report.

## The performance of the monitoring network for the previous year’s data (January 1 to December 31) is evaluated for the annual data certification package, which is due to U.S. EPA by May 1st. The AQS AMP600 Report, Certification Report, provides a summary of the performance of the network as to attaining the data quality objectives. The AMP600 report provides summary statistics on QA and QC activities and on data collected from each criteria pollutant monitor. The report also provides a summary evaluation of monitoring network performance by flagging data collection, and QA and QC activities as acceptable or green, warning or yellow, or recommend N or red. Normally, by this time any issues or concerns have already been resolved  and sufficient documentation is available. The annual certification letter provides a short summary to document data collection or QA and QC activities flagged as warning (yellow) or red (recommend N).

## Section 25:	References

- U.S. EPA method EQL-0895-107 (Flameless Atomic absorption)

- 40 CFR Part 50 Appendices L and M

- EPA QA/R-5 (March 2001), Requirements for Quality Assurance Project Plans, U.S. Environmental Protection Agency, Washington, DC

- EPA-454/B-18-006 (August 2018), Guide to Writing Quality Assurance Project Plans for Ambient Air Monitoring Networks, U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Air Quality Assessment Division, Research Triangle Park, NC

- EPA-454/B-17-001 (January 2017), Quality Assurance Handbook for Air Pollution Measurement Systems, Volume II: Ambient Air Quality Monitoring Program, U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Air Quality Assessment Division, Research Triangle Park, NC

- EPA QA/G-8 (November 2002), Guidance on Environmental Data Verification and Data Validation. U.S. Environmental Protection Agency, Washington, DC

- Quality Assurance Project Plan for the Federal Lead (Pb) Performance Evaluation Program, U.S. Environmental Protection Agency, Office of Air Quality Planning and Standards, Research Triangle Park, NC (EPA-457/B-14-001a, September 2014)

- EPA-454/B-16-001 (January 2016), Quality Assurance Guidance Document 2.12, Monitoring PM2.5 in Ambient Air Using Designated Reference or Class I Equivalent Methods

- QAPP: Calibration, Certification, and Verification Methods of Transfer Standards, Volume V (2022)