# Scope of Work (SOW): Real‑Time Temperature Monitoring for Vaccine Storage (VFC Program)

# Scope of Work

- Hardware: 30-60 monitoring nodes (one per refrigerator), each with an external, glycol‑buffered, NIST‑traceable calibrated probe; mounting hardware; power supplies; and onsite labels/asset tags. Include 10% spare probes and cables.

- Connectivity: Gateways and communication modules to support onsite network connectivity with automatic cellular failover. System must queue and forward data during connectivity disruptions without data loss.

- Software/Cloud Platform: Secure, role‑based web portal and mobile access with dashboards, alerts, audit logs, and reporting. Include unlimited user accounts for IDOH staff at no additional cost.

- Installation & Commissioning: Onsite installation, sensor placement, labeling, calibration certificate capture, and functional testing. Commission all 40 units with documented acceptance tests.

- Configuration: Alarm thresholds, notification rules, escalation trees (SMS, email, and voice/phone), and site‑specific schedules.

- Training & Change Management: Live training for end‑users and administrators (two sessions minimum, recorded), quick‑reference job aids, and SOP templates for alarm response.

- Maintenance & Support: Annual NIST‑traceable calibration (certificates provided), replacement of failed hardware, firmware/software updates, and 24×7 support with defined SLAs.

- Reporting & Data Retention: Automated weekly and monthly reports; on‑demand exports (CSV, PDF); retain data and audit logs for at least 7 years; IDOH owns all data.